| HGO1210/2016 Unique internal ID: 367517 | Sheet of legal act |
| --- | --- |
| The Republic of Moldova |  |
| GOUVERNMENT |  |
| DECREE No. 1210 of 03.11.2016 on the approval of the Sanitary Regulations on Radioprotection and Radiological Safety in the Practice of Nuclear Medicine |  |

Published: 11.11.2016 in the Official Gazette No. 388-398     art No: 1309     Date of entry into force: 11.12.2016

In order to implement the provisions of art.6 para. (2) and Art. 72 of the Law no.10-XVI of February 3, 2009 on state public health supervision (Official Gazette of the Republic of Moldova, 2009, no.67, art.183), with the subsequent modifications and completions, art.7 of the Law no. 132 of June 8, 2012 on the safe conduct of nuclear and radiological activities (Official Gazette of Republic of Moldova, 2012, no. 229-233, art.739), as well as in order to ensure a high level of protection of the general public's health, the Government DECIDES:

1. To approve the Sanitary Regulations on Radioprotection and Radiological Safety in the Practice of Nuclear Medicine (to be attached).

2. The Regulation and supervision of the safe conduct of the nuclear medicine practices are performed by the Ministry of Environment, through the National Agency for Regulation of Nuclear and Radiological Activities.

3. The state supervision of public health and exposure to ionizing radiation in nuclear medicine practices is performed by the Ministry of Health through the Public Health Supervision Service.

4. This Decree  shall enter into force in 30 days after the publication date.

PRIME MINISTER                                                 Pavel FILIP

Countersign:

Minister of Health                                                     Ruxanda Glavan

Minister of the Environment                                    Valeriu Munteanu

No. 1210. Chisinau, November 3, 2016.

Approved

through the Government Decree no. 1210

of November 3, 2015

SANITARY REGULATIONS

On ensuring of Radioprotection and Radiological Safety

in Nuclear Medicine Practices

The Sanitary Regulations on Radioprotection and Radiological Safety in Nuclear Medicine Practice (hereinafter - "Regulations") create the framework for the implementation of Directive 2013/59 / EURATOM of Council of Europe of 5 December 2013 of establishment of basic safety norms for the protection against presented hazards by exposure to ionizing radiation and repealing Directives 89/618 / Euratom, 90/641 / Euratom, 96/29 / Euratom, 97/43 / Euratom and 2003/122 / Euratom, Official Journal of the European Union (OJ) nr.L13/1   of 17 January 2014 and the application of Council Regulation No 1493/93 of 8 June 1993 on the carriage of radioactive substances between Member States, published in the Official Journal of the European Union (JO) L148 / 1 of 19 June 1993.

Chapter I

General dispositions

1. This Regulations establish specific requirements for ensuring the radioprotection of patients, staff, the population and the environment in the conduct of in vivo diagnostic investigations with the introduction of radiopharmaceuticals in the patient's body. The Regulations also establish the use of the set of reagents for radioimmunoassay in vitro for diagnosis purposes.

2. The Regulations aim to establish the requirements for ensuring radiological safety and radioprotection of patients, preventing the exposure of exposed workers and the general public to ionizing radiation, the preventing of contamination of the environment by conducting radiopharmaceutical investigations.

3. The provisions of these Regulations shall apply to:

1) the departments, the nuclear medicine laboratories of the medical-sanitary institutions, the scientific research institutions, the universities, the vocational and medical institutions, where the methods of nuclear medicine are used;

2) the regulatory body, supervision and control body in the field of nuclear and radiological activities, state authorities and institutions with responsibilities in the field of nuclear and radiological activities, institutions using open ionizing irradiation sources, applied in the field of nuclear medicine, state bodies and institutions tasks of supervision and control in the field of nuclear and radiological activities, design institutions, central and local public authorities, organizations and enterprises, irrespective of the property type and organizational-legal form;

3) To the institutions, that perform estimation, recording and control of patient exposure and exposure to ionizing radiation.

4. The provisions of these Regulations shall apply to the phases of design, location, construction, reconstruction and operation of the nuclear medicine premises, laboratories and facilities.

5. The used terms and expressions in these Regulations are defined in Law no.10 of 3 February 2009 on State Public Health Supervision and Law No 133 of 8 June 2012 on the Safe Operation of Nuclear and Radiological Activities.

6. For the purposes of these Regulations, the used terms shall have the following meanings:

quality assurance in nuclear medicine - a set of planned and performed measures by the radiological authorization holder, the purpose of which is to obtain the expected results of the radiopharmaceutical investigations in conditions of radiological safety, radioprotection of the exposed workers and of the patients, meant to guarantee a degree corresponding to quality in a structure, system or component;

room for deactivation - intended space for the decontamination of utensils, medical instruments and elements used in the work with radiopharmaceuticals;

procedural room - intended space for the introduction of radiopharmaceutical preparations into the patient's body;

radiometric room - space for radiometry of the organs, tissues and body of the patient;

preparation room - intended space for performing the technological procedures, preparation, packaging and preparation of patient radio-pharmaceutical products, including procedures with radionuclide generators;

PET center - specialized diagnosis subdivision with in vivo designed radionuclides for performing PET diagnostics and the production of positron emitting radiopharmaceutical preparations with quality control;

quality control - a system of organizational measures, technical means and technological procedures for the quantification, monitoring and maintenance at an optimal level of the characteristics of the radiodiagnostic equipment, the diagnostic regimes, as well as the quality parameters of the radiopharmaceutical products;

diagnosis with in vivo radionuclides - establishing the existence, nature and spread of the pathological process in the patient's body, that is based on imaging visualization and / or determining the spatial and temporal distribution of the radiopharmaceutical preparations , that are introduced into the patient's body;

radioactive waste repository - all intended rooms for the temporary storage of liquid and solid radioactive waste for the purpose of its exposure to disintegration and / or accumulation for subsequent centralized storage;

storage of radiopharmaceutical preparations - space for the temporary storage of radiopharmaceuticals, that are ready for entry and prepared in the radiopharmaceutical laboratory for in vivo diagnosis purposes;

gamma room - a stationary or mobile installation for scintigraphy, including a gamma-ray scintillation detector, stands assembly, patient bed, an electronic signal conversion system, and a computer for forming and viewing scintigraphic images;

gamma tomography - stationary computer tomography installation with gamma photon emission, that is containing one or more gamma-sensitive detectors, a tripod with detector rotation mechanism around the longitudinal axis, a patient's bed, an electronic detector signal conversion system and a computer for reconstruction and viewing tomography images;

radionuclide generator - portable device with local radiological protection for the rapid acquisition of short life radionuclides under the conditions of the medical institution;

stationary means of radioprotection - elements, construction structures and protective equipment, that are intended for the protection of professional exposures and patients against external radiation, inhalation or intrusion of radioactive substances into the body;

individual means of radioprotection - the used technical means by professional exposures to conduct radiodiagnostic examinations for the protection of the body or of certain organs during the investigations;

radiopharmaceutical preparations - preparations and pharmaceutical compounds with radionuclides, registered by the Ministry of Health and authorized according to art. 24 and art. 43 paragraph (2) of the Law no.10-XVI of 3 February 2009 on the state public health supervision for use in the diagnosis of various diseases;

nuclear medicine procedures - diagnostic procedures with the introduction of radiopharmaceuticals into the patient's body;

scintigraphy - performed diagnostic procedures at the gamma room for viewing images, projection, spatial distribution of radiopharmaceuticals in the patient's body (static scintigraphy) or recording of time characteristics of the stationary (space-time) distribution of radiopharmaceutical preparations in the patient's body;

positron emission tomography PET - (Positron Emission Tomography) - is the abbreviation of tomography on positron emission basis . PET is one of the most modern methods of obtaining functional images for medical purposes. It is a diagnostic procedure to visualize the spatial distribution of the positron emitting radiopharmaceutical preparation into the patient's body by annihilation irradiation.

7. The radioprotection system at the design and operation stages of the nuclear medicine sections will provide for the implementation in practice of the following basic radiation protection principles - normalization, justification and optimization.

8. The normative principle is to establish norms (limits of permitted irradiation doses, admissible levels and levels of control) of ionizing radiation exposure of professionally exposed workers and of the general public.

1) The effective dose limit for occupational exposure is 20 mSv for each year. However, in special situations or in certain exposure situations, a higher effective dose of up to 50 mSv may be authorized by the competent authority in one year, provided that the average annual dose for any five consecutive years, including the years for which the limit has been exceeded, shall not exceed 20 mSv.

2) In addition to the actual dose limits set out in subsection 1), the following equivalent dose limits apply:

a) the equivalent dose limit for crystalline is 20 mSv in one year or 100 mSv for any period of five consecutive years subject to a maximum dose of 50 mSv in one year;

b) the equivalent dose limit for the skin is 500 mSv per year; this limit applies to the average dose for any area of ​​1 cm2, irrespective of the exposed surface;

c) the equivalent dose limit for the extremities is 500 mSv per year.

3) For public exposure, the effective dose limit is 1 mSv per year.

4) In addition to the dose limit provided in subp. 3), the following equivalent dose limits apply:

a) the equivalent dose limit for the lens is 15 mSv per year;

b) The equivalent skin dose limit is 50 mSv per year, calculated as an average for any surface of ​​1 cm2, regardless of the exposed surface.

5) The radiological authorization holder shall ensure that the effective dose limit for disciples of  age 16-18 and for students of age 16 to 18 who are required to work with radiation sources in nuclear medicine laboratories during their studies is 6 mSv per year.

6) In addition to the effective dose limit, provided in subp. 5), the following equivalent dose limits apply:

a) the equivalent dose limit for the crystalline is 15 mSv per year;

b) the equivalent skin dose limit is 150 mSv per year, calculated as an average for any area of ​​1 cm2, regardless of the exposed area;

c) the equivalent dose limit for the extremities is 150 mSv per year.

7) The holder of the radiological authorization shall ensure that the protection of the fetus is comparable to that provided to the general public. As soon as a pregnant worker informs the undertaking or, in the case of an external worker, the employer, on the pregnancy, the employer and the employer ensure that the working conditions for the pregnant worker are such that the equivalent dose to which the fetus is exposed, is less possible and does not exceed 1 mSv at least for the remaining period of pregnancy.

9. The working conditions of professionally exposed pregnant women should ensure that the received effective dose by the fetus does not exceed 1 mSv. From the moment a woman announces the administration of the medical institution that she is breastfeeding, she is no longer used in activities, involving a high risk of radioactive contamination and exposure to ionizing radiation. The administration of the medical institution must provide a pregnant woman with a place of work that is not related to the use of ionizing irradiation sources during the pregnancy and breastfeeding of the child from the day receiving the information about pregnancy.

10. The principle of normalization does not apply to patients in the conduct of diagnostic investigations by the in vivo method with the introduction of radiopharmaceuticals. This principle is achieved by establishing exposure norms (admissible dose limits) for occupational exposure and population.

11. When conducting diagnostic investigations using the in vivo method with the introduction of radiopharmaceuticals, the justification principle is achieved by fulfilling the following requirements:

1) performing diagnostic investigations with the introduction of radiopharmaceuticals only in cases where the obtained diagnostic information (benefit) for the patient is greater than the damage caused to the health by the action of the ionizing radiation;

2) conducting diagnostic investigations with radiopharmaceuticals only according to the medical (clinical) indications in cases where there are no other alternative methods of diagnosis or when they cannot be performed or are insufficiently informative (priority and initial use of alternative non-ionizing methods);

3) choosing the most harmless methods for radiopharmaceutical examinations;

4) use of approved radiopharmaceutical investigation methods;

5) in vivo diagnosis only of those authorized radiopharmaceuticals by the Ministry of Health;

6) ensuring the necessary conditions for obtaining reliable diagnostic information;

7) all new types of practices involving exposure to medical purposes are justified before being generally adopted;

8) Existing types of practice, involving exposure to medical purposes can be analyzed whenever new and important evidence emerges about their effectiveness or consequences;

9) All individual medical exposures for medical purposes are warranted in advance, taking into account specific exposure objectives and individual patient characteristics.

If a type of practice, involving medical exposure is not generally justified, a specific individual exposure of this type could be justified only for the specific case.

The medical practitioner and the curative doctor seek, as far as possible, to obtain previous diagnostic information or medical records relevant to the planned exposure and take these data into account in order to avoid unnecessary exposure;

10) medical exposure for biomedical and medical research is reviewed by the National Ethics Committee for Clinical Study;

11) special consideration shall be given to justifying exposures for medical purposes where there is no direct benefit to the health of the exposed person.

12. The principle of optimization in conducting diagnostic investigations with the introduction of radiopharmaceuticals consists in obtaining the necessary diagnostic information, qualitative with radiation exposure at the lowest possible reasonable level, taking into account the economic and social factors - ALARA principle (As Low as Reasonable Achievable - as low as reasonable achievable), established by the recommendations of the International Atomic Energy Agency and other relevant international authorities.

13. When monitoring and supervising workplaces and working conditions, professionally exposed workers to ionizing radiation shall be divided into:

1) workers who are professionally exposed Category A - persons who work with sources of ionizing radiation in nuclear medicine, including persons performing radiological control. These workers may receive an effective dose of more than 5 mSv / year or an equivalent dose greater than 3/10 of the limit values ​​for crystalline, skin and extremities;

2) Category B exposed workers - persons exposed to ionizing radiation, that do not fall into Category A and receive effective or equivalent doses less than those assigned to Category A personnel.

14. Lists of professionally exposed workers shall be approved by the administration of the medical institution and shall be coordinated with the State Health Supervision Service of the Ministry of Health on the basis of art. 30 pt. 6, art.42 par. 3, art.43 par. (2), (3) and (4) of the Law no.10-XVI of 3 February 2009 on state public health supervision.

15. For patients who are not assigned to categories A and B, patients who are not investigated with the use of radiopharmaceuticals, visitors, relatives of patients and other persons, who are not assigned to occupational exposure to radiation, the annual exposure doses will not exceed the annual dose limits set for the public.

16. The estimation of the existing system for ensuring radiological safety and radioprotection in nuclear medicine departments is based on the following indices:

1) the characteristic of possible and actual radioactive contamination of the environment;

2) the analysis of the measures for ensuring the radiological safety and radiation protection and the observance of the provisions of the normative documents in the field of radioprotection;

3) the probability of occurrence of incidents and accidents with radiation and possible consequences;

4) the medical training of the medical institution towards the elimination of radiation incidents and accidents and their consequences;

5) analysis of received exposure doses by exposed workers to ionizing radiation and members of the public professionally during the activities of nuclear medicine units;

6) the number of exposed people in values ​​exceeding the established dose limits.

17. Radiological security and radioprotection in the nuclear medicine subdivision shall be ensured by:

1) the quality of the design of the rooms, in which the nuclear medicine subdivision is located;

2) justified selection of the site and land for the location of the nuclear medicine unit block;

3) justified operation of nuclear machinery and equipment;

4) physical protection of radioactive sources;

5) fine-tuning of the established technologies for radiopharmaceutical investigations;

6) the permanent presence and functioning of the quality assurance and quality control system;

7) Planning and performing specific concrete activities on ensuring radiological safety, radioprotection of patients, professionally exposed workers and the general public in the daily activity program of the section, at the stages of reconstruction and cessation of exploitation;

8) increasing the knowledge of patients and the public in the field of radioprotection.

18. Diagnostic investigations with the use of radiopharmaceuticals should be performed only in accordance with clinical prescriptions of the clinician in the presence of clinical indications.

19. The optimal activity of the radiopharmaceutical preparation is ensured in the patient's body to ensure reliable diagnostic information at the lowest possible dose - according to the ALARA principle. The patient has to comply with the requirements and established recommendations by the practitioner.

20. Radiation protection of personnel in the nuclear medicine subdivision shall be ensured by:

1) limiting access to activities with sources of ionizing radiation according to age, sex, health status, previous occupational exposure and accidental exposure;

2) the presence in sufficient quantities of the collective radiation protection means and the limitation of the duration of activities with radioactive sources;

3) strict observance of the requirements and recommendations for personnel, which are set out in the service instructions and the normative-technical documentation;

4) use of individual radiation protection means;

5) Effective measures to prevent and eliminate radiological accidents.

21. The radioprotection of persons, who are not assigned to the professionally exposed personnel but accompanies the patient assists them during the procedures and cares for the patient, who has been given radiopharmaceuticals is ensured by observing the requirements and recommendations for the public which are described in the instructions service on nuclear medicine activities in this medical institution.

22. Diagnostic investigations with radiopharmaceuticals are performed only in medical institutions, which have specialized nuclear medicine subdivisions, equipped with the necessary equipment, supplemented with professional training staff in the field of activity with open sources of ionizing radiation.

23. The use of radiopharmaceuticals for clinical application is also prohibited, as well as the application of in vivo diagnostic methods to their introduction into the patient's body unless they have been approved and authorized by the Ministry of Health in accordance with Article 24 and Article 43 paragraph (2) of the Law no.10-XVI of 3 February 2009 on state public health supervision.

24. The design and construction of the rooms of the nuclear medicine subdivisions, the manufacture and servicing of the equipment, the technological equipment and the means of radioprotection shall be performed only by the institutions holding the radiological authorization and the sanitary approval for the execution of these activities.

25. The authorization of nuclear medicine activities, including the radiological authorization and the sanitary authorization of the functioning of the nuclear medicine units, shall be performed on the basis of the submitted requests by the user of the ionizing radiation sources, based on the evaluation of the submitted documents. The set of submitted documents must include:

1) the project of the nuclear medicine unit, the project of the installation of nuclear medicine installations and equipment;

2) calculations of stationary protection;

3) staff and patient radiation protection program;

4) the program for monitoring the exposure to ionizing radiation and the estimation of the health of the employees;

5) data on the professional training of the personnel.

26. Sanitary endorsement are subject to the following activities, draft documents, products and services in the frame of nuclear medicine practices:

1) assignment of land for construction and reconstruction;

2) construction and reconstruction projects;

3) final reception of constructed and reconstructed buildings;

4) standards for products and services;

5) production technologies.

27. State public health supervision in nuclear medicine departments includes estimating working conditions and harmful factors, monitoring exposure to ionizing radiation, and medical supervision of the health status of professionally exposed workers and patients.

28. When conducting radiopharmaceutical investigations, holders of radiological authorization shall provide:

1) control of the validity of authorization documents (radiological authorization, certificate of security, sanitary authorization);

2) control of the maintenance, servicing and repair conditions of the plants acting on the safety of the practiced radiological activities, to a sufficient extent to ensure their proper functional status. The used installations and equipment in nuclear medicine practice will have the necessary operational documentation (valid certificates of security);

3) access of representatives of the National Agency for Regulation of Nuclear and Radiological Activities (hereinafter - Agency) and other authorities with powers of supervision and control in the field of nuclear and radiological activities, according to the established competences by the Law no.132 of 8 June 2012 on the deployment in the safety of nuclear and radiological activities for the purpose of performing inspections at the objectives, where authorized activities are performed with ionizing radiation sources;

4) elaboration and realization of the action plan for the radioprotection and radiological safety system;

5) knowledge and observance of normative acts regarding radioprotection and radiological safety and full compliance with their provisions;

6) fulfillment of the requirements of the normative documents regarding the radioprotection and radiological safety under the normal operation of the radiological installations, as well as in case of accident;

7) periodical analysis and estimation of radioprotection, radiological safety in the medical institution, carrying out measures for their improvement according to an annual approved plan by the authorization holders;

8) inventory, control and inventory of radioactive sources and wastes according to the approved registers;

9) temporary storage of radioactive sources, the continued use of which is not required (including radioactive sources with expired use) by specialized organizations for the storage and treatment of radioactive waste;

10) exploitation of radioactive sources only during the indicated period of operation;

11) the physical protection of radioactive sources in order to avoid theft and unauthorized use;

12) the required number and qualification of the employees (workers exposed professionally) who are involved in authorized activities;

13) planning and carrying out the improvement of employees (professionally exposed workers) in the field of radioprotection and radiological safety, radiological monitoring, recording and control of radioactive sources and radioactive waste, ensuring their physical protection;

14) monitoring exposure to ionizing radiation at workplaces, rooms and on the territory of the medical institution;

15) individual dosimetric monitoring of exposed workers to occupational exposure;

16) periodical training, perfection and attestation of the managers of the nuclear medicine units, radioprotection specialists, radioprotection specialists, as well as of other persons permanently or temporarily performing radiopharmaceuticals in the field of radioprotection and radiological safety;

17) regular training and examination of the knowledge of professionally exposed workers to radioprotection and radiological safety;

18) performing preventive and periodic medical examinations of professionally exposed workers;

19) permanent information of exposed workers to occupational exposure levels at the workplace and the received individual exposure doses;

20) information on the basic aspects of ionizing radiation, their effects and the actions to be taken at the place of the radiological incidents (detection or suspected presence of a source);

21) timely notification of the Agency, the Public Health Service of Public Health and other bodies with responsibilities in the field of nuclear and radiological activities, as appropriate, on radiological incidents and accidents;

22) fulfilling the requirements and prescriptions of the Agency, the bodies with responsibilities in the field of nuclear and radiological activities;

23) realization and dissemination of citizens' rights in the field of radioprotection and radiological safety;

24) control of the elements and conditions included in the radiological authorization;

25) use of verified metrological measuring equipment and instruments.

29. The responsibility for ensuring radioprotection at workplaces, workplaces in nuclear medicine departments is assigned to the administration or management of the medical institution. The responsibilities of the radiological authorization holder are described in detail in Chapter VII of these Regulations.

30. The administration of the medical institution shall elaborate and approve the instructions for radioprotection and radiological safety, the requirements for activities with ionizing radiation sources, the registration, storage, transport, teaching of radioactive sources, including their radiological control, provisions on personal hygiene measures and work hygiene, as well as the actions of professionally exposed workers to incidents and accidents with radiation.

31. During radiological activity in nuclear medicine units, a negative impact on exposed workers to the workforce may have the following harmful and dangerous factors:

1) ionizing radiation:

a) increased external-gamma radiation in the service rooms of the department, the laboratory;

b) presence of radioactive contamination of working surfaces, increased concentrations of aerosols and radioactive gases in the air of the service rooms;

c) internal exposure through the penetration of radionuclides into the body;

2) non-radioactive noxious factors in the occupational environment:

a) high voltage values in the circuitry of the equipment and the equipment, the short circuit through the human body;

b)  high level of noise pollution, produced by the used equipment, the refrigeration plant, the fans and the facilities, the increased physical effort during the operation of the equipment;

c)  influence of harmful chemicals used for the synthesis or preparation of radiopharmaceuticals in the operation of apparatus and equipment as well as in the deactivation and room cleaning in the nuclear medicine departments.

32. The levels of harmful and dangerous factors in the nuclear medicine sections shall not exceed the admissible established values by the radiation protection regulations.

CHAPTER II

MEDICAL EXPOSURE. JUSTIFICATION AND

OPTIMIZATION OF MEDICAL EXPOSURES

33. The holder of the radiological authorization shall ensure that:

1) the patients are only exposed, if exposure has been prescribed by a practitioner;

2) the specialists in nuclear medicine to ensure the radioprotection of patients, both by prescribing and by the way of administration of the dose;

3)  the exposure of persons, that are supporting or caring for patients and family members of patients is subject to dose constraints;

4) the quality control requirements are fulfilled in the case of diagnosis with open sources by an accredited institution in the national accreditation system by the National Accreditation Center;

5) the personnel are trained in the use of radioactive sources and used equipment in nuclear medicine, equipment for the measurement and detection of radiation, safety systems and devices, appropriate to their function and responsibility, and know the applicable radiological emergency;

6) the personnel comply with applicable rules and procedures for radioprotection of patients.

34. The medical exposures must be justified, taking into account the benefits of diagnosis, the risk of exposure to radiation and the existence of alternative methods not involving exposure to ionizing radiation: ultrasound, nuclear magnetic resonance.

35. The exposures of persons, who participate voluntarily in medical or diagnostic biomedical programs are justified only if the benefit is net superior to detriment, that ionizing radiation brings to the individual.

36. The exposures of persons, that are participating in medical or biomedical investigation programs shall be made in accordance with the norms provisions of the radioprotection in force. The holder of the radiological authorization shall ensure that the nuclear physician, who prescribes or drives the diagnosis exposure:

a) keep the exposure to the minimum necessary to obtain the desired information;

b) consider all relevant information from previous exposures in order to avoid any further exposure;

c) consider the reference levels for medical exposures set out in radioprotection norms in force for the radioprotection of persons in the case of medical exposures.

37. The involved personnel in medical exposure reduces patient exposure without diminishing image quality by:

a) selecting the most appropriate radiopharmaceutical and its activity;

b) use of methods to block the absorption of the radiopharmaceutical in the non-study organs and, respectively, to accelerate excretion, as appropriate;

c) use of appropriate procurement and processing methods;

d) avoiding the administration of radiopharmaceuticals to pregnant women or who might be pregnant, except when the life and health of the pregnant woman is at risk;

e) cessation of breastfeeding for nursing women until the radiopharmaceutical is no longer excreted in milk in expected amounts to result in an unacceptable effective dose to the newborn;

f) radiopharmaceuticals administration to children for diagnostic purposes only if there are clear clinical indications, and the administered radiological activity is reduced proportionally by weight, body surface.

38. The holder of the radiological authorization must inform in writing:

1) patients with the associated risk with the procedure;

2) patients who have received radiopharmaceuticals to minimize contact with family members, minors and women, who may be pregnant or are pregnant.

CHAPTER III

REQUIREMENTS FOR THE LOCATION AND

FACILITATION OF NUCLEAR MEDICINE LABORATORIES

39. Nuclear Medicine Laboratories are going to be located in a separate block or wing fully isolated from the rest of the building, on the ground floor and with a separate entrance. The placement of nuclear medicine laboratories in residential buildings and children's institutions are prohibited. It is not allowed to be located in the neighborhood of the radioisotopic diagnostic laboratory, where the works are performed with I and II danger degrees (horizontally and vertically) of the salons for children and pregnant women.

40. Nuclear medicine laboratories for in-vivo radiodiagnostic with radioactive sources are going to be located in a separate block or in an isolated wing from the rest of the rooms, and those for in-vitro radiodiagnostic use with radioactive sources do not require special location conditions. The establishment of the sanitary protection area from the nuclear medicine laboratories is not foreseen.

41. At the placement of nuclear medicine laboratories within buildings, where other activities are performed, the rooms, where the laboratory-specific activities are performed will be located in one part of the building, separated from other non-nuclear activities in such a way as to ensure the circuit functional of laboratory.

42. In laboratories, using radioactive sources for in vivo diagnosis, control units for heating, natural gas supply, water supply and electrical installations shall be located outside the premises, where the activity is performed.

43. At the simultaneous use in the nuclear medicine laboratory of open in vitro and in vivo diagnostic radioactive sources, the specific workspaces to each type of arrangement shall be clearly separated and delimited.

44. The factors, that need to be taken into account for the establishment of the nuclear medicine laboratory design category are:

1) the annual limit of incorporation for radionuclides, that are going to be used, provided in the radioprotection norms in force, the type of radionuclides and the performed operations;

2) the risk of external exposure and contamination, associated with the types of involved radionuclides;

3) type of diagnostic technique;

4) maximum daily, weekly, monthly or annual workload.

45. The Nuclear Medicine Laboratory's design categories represent the degree of protection allowed and are:

1) Category 1: nuclear medicine laboratories for in vivo diagnosis;

2) Category 2: nuclear medicine laboratories for in vitro diagnosis.

46. ​​In the case of nuclear medicine laboratories, that are using radioactive sources for in vitro / in vivo diagnostics, flooring, carpentry and work surfaces in rooms must be made in such a way that they can be easily washed and decontaminated. In this respect, the following conditions shall be met:

1) the floor must be covered with non-absorbent materials, that can be cleaned and cleaned easily, resistant to decontamination of the agent’s actions;

2) the floor must withstand the weight of the screens and be smooth, free of cracks and joints, and the connection between the floor and the walls must be sealed and rounded;

3) the walls and ceilings shall be made of any kind of profile or ornamentation, with all angles and rounded edges, and shall be covered with non-absorbent, easily washable and resistant to decontamination of the agent’s actions;

4) for doors and windows are recommended the metallic joinery or PVC with the simplest profiles; the windows will be fixed and the doors will be flat, preferably swinging;

5) The work surfaces must withstand the weight of the screens and be covered with non-absorbent, easily washable and resistant to the chemical agents, with which they work. These surfaces will be covered with thin sheets of absorbent materials, which after use will be considered radioactive waste.

47. The furniture used in nuclear medicine laboratories should be as simple as possible, have a smooth, non-absorbent surface and be covered with easily decontaminated materials.

48. Nuclear medicine laboratories using radioactive sources for in vivo diagnosis shall be provided with:

1) equipped rooms with niches, boxes or other ventilated and shielded enclosures for the preparation of radiopharmaceuticals and dilutions of radioactive substances;

2) room for the storage of radioactive sources and / or radioactive waste;

3) room for injecting patients or administering radiopharmaceuticals;

4) room for conducting medical diagnosis in vivo;

5) consulting room for practitioner;

6)separate waiting rooms for patients, being administrated with radiopharmaceuticals and non-radiopharmaceuticals;

7) personnel showers installations;

8) dosimetric control point at the exit of the controlled areas;

9) space for the decontamination and sterilization of the technical-medical instrument;

10) shielded containers for the temporary collecting and storage of radioactive waste, arising from the practice;

11) health care group for patients receiving radiopharmaceuticals, separated from the health care team or patients, who have not been given the radiopharmaceuticals.

49. Nuclear medicine laboratories using radioactive sources for in vitro diagnosis shall be provided with:

1) room for storage and preparation of solutions;

2) niches, if there are works, where air contamination is possible;

3) shower facilities;

4) room for the storage of radioactive sources and / or radioactive waste;

5) refrigeration space in the radioactive source deposit;

6) room for the processing of radioactive biological products;

7) space for conducting in vitro tests;

8) space for in vitro radiometry of biological products;

9) space for the decontamination and sterilization of the technical-medical instrument;

10) space for the harvesting of biological products.

50. If diagnostic testing with radiopharmaceuticals requires testing on laboratory animals, the functional circuit of that area will be completely separate from areas for human medical activity.

51. The composition and the surfaces of the rooms of the nuclear medicine laboratories are indicated in Annex 1 to these Regulations.

CHAPTER IV

REQUIREMENTS RELATING TO PLACING,

PLANNING AND FITTINGS OF PET CENTER LOADS

52. Designing the center's rooms with PET technology and its technical endowment is going to be performed by a specialized licensed institution, that is holding the radiological authorization in the field, taking into account the complexity and requirements of the PET technology supplier (manufacturer).

53. The PET center is equipped with separate rooms for staff and patients as well, with a passage for hospitalized patients with cold-weather heating. The PET-center halls must not be passable for foreigners.

54. The radiochemical complex of cyclotron is a separate controlled area with restricted access for foreign persons, including patients, with a separate entrance, provided with an air lock, emergency exit and certified rooms for the production of radiopharmaceuticals.

55. The cyclotron shall be placed in a special canyon with concrete and ceiling walls, the thickness of which shall be calculated in accordance with the requirements of the radiological protection requirements and the requirements of the PET technology supplier (manufacturer).

56. The PET-center compartments with the radiochemical complex of the cyclotron are placed on the ground floor of the building or in a separate block. It is allowed to place the cyclotron lower than the ground floor of the building in compliance with the radioprotection requirements and the building norms.

57. The cyclotron and machinery assembling technology will provide for a minimum clearance of 2.8 x 2.8 m in the walls of the cyclotron canopy and the radiochemical laboratory or in the ceiling of these rooms. For proper-radioprotected cyclotroanes, these opening spaces must be larger and meet the requirements of the cyclotron-producing company and the radiochemical equipment.

58. The PET-center is provided with paved roads for access to trucks and cranes.

59. Entry into the cyclotron bunker shall be provided with a protective door fitted with at least four locking systems in accordance with the requirements and recommendations of the cyclotron-producing company.

60. The cyclotron bunker, the radiochemical laboratory, and the diagnostic center of the PET center are located adjacent to or at a minimum distance from each other. The distance from the cyclotron canyon to the radiochemical laboratory should not exceed 20 m. The mentioned rooms are to be joined with hidden channels for the location of gas pipes, water and other communications in accordance with the requirements and recommendations of the manufacturing company.

61. A separate room for gas balloon placement is provided, connected with the corresponding lines of the pipelines with the cyclotron canyon and the radiochemical laboratory.

62. Depending on the range and quantities of synthesized radiopharmaceuticals in the radiochemical laboratory, one or more "hot" boxes and / or several mini-boxes for PET radiochemistry will be installed in accordance with the requirements and recommendations of the manufacturing company.

63. The temporary storage deposit for radiopharmaceuticals and the procedure room shall be equipped with reflux cabinets, inside which shall be located protective barriers.

64. The temporary storage deposit for radioactive waste shall be provided with protective walls, where plastic bags or closed containers for radioactive waste held for disintegration shall be kept, also a hermetically sealed metal closet with a local pump for non-radioactive waste from solutions technology.

65. A sanitary filter is placed near the external entrance to the radiochemical complex of the cyclotron (considering the extent of the possible production of I-123 to the cyclotron) and the other between the radionuclide supply block and the PET-center general block. The floor, the walls and the ceiling of the sanitary filter for patients are made of moisture-resistant materials with poor absorption properties of the radioactivity, which can be cleaned and cleaned easily, resistant to the action of decontamination agents.

66. Patient and personnel WC are designed separately with disabled access for patients in the personnel's WC. The patient's WC is located in the immediate proximity of the procedure room and the diagnostic center of the PET-center have to be equipped with a water guide and washing installation, equipped with a special pedal.

67. The requirements for the dimensions, configuration and technical assurance of the cyclotron canyon, the radiochemical laboratory, the positron tomography cabinet are determined by design, according to the technical passport of the concrete installations, based on the recommendations of the apparatus of the manufacturers and equipment and are coordinated with Agency and Supervision State Service of Public Health.

68. In the cyclotron canyon are provided the radiochemical laboratory, and the positron tomography cabinet, special ventilation and air conditioning systems, which work separately with reflux-ventilation ventilation of other PET-center rooms.

69. The ventilation of the cyclotron canyon, radiochemical laboratory, and positron tomography cabinet is autonomous with air-reflux-flow in the upper zone and removal of the area from two upper-2/3-volume areas and 1/3 at the bottom.

70. The reflux aero conduits in said rooms are equipped with air purifiers with filters and adsorbents of radioactive gases. The ventilation is about to work permanently.

71. The need for additional ventilation facilities and the air exchange rate is determined by the project based on the recommendations of the manufacturing company.

72. The air removed from the rooms, the boxes and the evacuation cabinets is pre-cleaned by means of the own filters, provided by the construction elements. In rooms and boxes with closed apertures shall be provided a deficit of minimum of 20 mm of water column. The cameras and boxes are equipped with speed controllers. Radioactive gas filters and adsorbents are installed in the immediate vicinity of the chambers, boxes and exhaust cabinets to minimize radioactive contamination of the air ducts.

73. The ventilation of the general rooms of the PET-center is projected by the influx-reflux type with the removal and supply of air in the upper area of the room with the air exchange rate at the influx-3 at the reflux-4.

74. The residual waters in the premises of the radionuclide supply block (except for the radiopharmaceutical synthesis room) and the general building block are allowed to be discharged into the household waste.

75. The technological solutions that have been used to wash and disable the boxes in the PET center are kept in glass bottles that are installed in the hermetic cupboard with a local pump.

76. The composition and surfaces of the PET-center rooms are shown in Annex 2 to these Regulations.

CHAPTER V

REQUIREMENTS FOR VENTILATION,

AIR PURIFICATION AND HEATING

77. The ventilation and air purification systems in nuclear medicine laboratories shall be designed to prevent contamination, including radioactive contamination of air in working rooms and atmospheric air, from exceeding the maximum permissible limits. The number of suction ventilation systems is to be minimized.

78. The movement of air between areas is going to be done from the unattended area to the supervised area and then to the controlled area. The pressure in the controlled area is to be smaller than the other areas.

79. The introduction and suction holes shall be provided with suitable filters and positioned so as to ensure effective removal of the air from the rooms, to ensure easy replacement of the filtering agent.

80. In buildings where only a portion is used by the nuclear medicine laboratory, the ventilation system for rooms where open radioactive sources are operated are completely separated from the system intended for other rooms where radioactive sources are not being used.

81. Before being discharged into the atmosphere, contaminated air in rooms, boxes, niches will be passed through located filters as close as possible to sources of pollution to minimize contamination of the main air ducts. The used filters will be treated as radioactive waste. Appropriate monitoring of exhaust gas effluents will be ensured.

82. The nuclear medicine laboratories, that are using radioactive sources for in vivo diagnosis outside of air purification systems shall be fitted with escape hatches the height of which is intended to ensure adequate air dispersal.

83. The air recirculation and room ventilation shall be prohibited without the use of mechanical filtration systems in nuclear medicine laboratories in which radioactive sources for in vivo diagnosis are used.

84. The air-entraining fans shall be positioned such that the connected fresh air ducts to the suction side of the fans do not pass through contaminated rooms or spaces, and air suction fans in contaminated rooms shall be so placed that the connected ducts to the fan outlet of the fans should not pass through rooms or spaces, whose contamination should be avoided. The admitted atmospheric air will cool down to about 200 ° C.

CHAPTER VI

REQUIREMENTS FOR WATER

SUPPLY AND SEWERAGE

85. The nuclear medicine laboratories are supplied with hot and cold water supply systems, connected to public networks.

86. In nuclear medicine laboratories for which radioactive sources for in vivo diagnosis are used, apart from the normal sewerage system and separate from it, a radioactive sewerage system with two distinct sewers shall be provided for:

1) radioactive effluents - exceeding the release limits in accordance with the provisions of regulations in force on the management of radioactive waste, which must be directed to a treatment plant or to storage tanks, from which they can subsequently be released into the environment;

2) radioactive suspected effluents, that are going to be directed to retention reservoirs and, after prior radioactivity control, discharged to public sewers or if necessary treated as liquid radioactive waste.

87. For nuclear medicine laboratories using radioactive sources for in vitro diagnosis, it is necessary to provide for a radioactive sewerage system for suspected radioactive effluents, apart from normal sewerage and separate from it. Sewerage systems and tanks must be protected against corrosion.

88. The tanks for the collection of radioactive effluents shall be easily accessible and secured against leakages of radioactive fluid. Tanks shall be provided with sampling facilities for the periodic analysis of radioactive contents.

CHAPTER VII

RESPONSIBILITIES

89. The holder of a radiological authorization has the following obligations:

1) to ensure the conduct of the practice of nuclear medicine in accordance with the provisions of the Law no.132 of 8 June 2012 on the safe conduct of nuclear and radiological activities and of Law no.10 of 3 February 2009 on state public health supervision;

2) to notify the Agency and the State Health Supervision Service of any intention to make changes to the authorized practice;

3) to provide adequate spaces and arrangements for the conduct of nuclear medicine;

4) to ensure appropriate equipment, dosimetry control equipment, appropriate work equipment and radioprotection equipment, accessories necessary to perform the practice;

5) to provide the necessary conditions for all used installations, apparatus and equipment to be in good working order and correctly maintained;

6) to provide professional A-class exhibitions with authorized dosimetric supervision means and services, accredited dosimetry laboratories;

7) to provide the means for radiological monitoring and the working environment and maintain a record of the results;

8) to ensure, that the dosimetric control apparatus is in a functional and verified metrological condition;

9) to prepare and implement an operational radiological protection system, including the applicable procedures and standards to the safe use of radiological installations and radiation sources by operators;

10) to designate radiation and radiological safety officers for each controlled area in order to ensure the implementation of the operational radiological protection system;

11) to consult or hire medical physicians;

12) to provide for radiological safety assessments to identify normal exposure and predictable exposure sources, to estimate the probability and magnitude of the resulting doses, and to assess the means and methods necessary to provide radioprotection and radiological safety;

13) to ensure that involved personnel in the practice of nuclear medicine have the necessary training and exercise permit, to ensure the necessary conditions for their participation in the training course;

14) to provide the necessary resources for the appropriate training of the involved personnel in the implementation of the practice in the following aspects: radioprotection and radiological safety, the working procedures, the regulations in the field and the incident response plan;

15) to organize the recording of the results of individual monitoring for each professionally exposed person to category A in the course of the practice of nuclear medicine, as well as for all exposed persons to accidental or emergency exposure;

16) to provide the necessary resources to implement the contingency plan in the event of a radiological incident;

17) to provide the necessary equipment for intervention in case of incidents, involving the used sources of ionizing radiation;

18) to ensure the provided conditions in specific regulations for disposing of radioactive waste, returning to the supplier / manufacturer and / or handing over radioactive waste, arising from the practice of nuclear medicine to authorized organizations;

19) to ensure the medical supervision of the professionally exposed personnel according to the Regulations of the Ministry of Health and special medical supervision of the exposed persons, who have been exposed at doses above the dose limits provided by the radiation protection norms in force;

20) to ensure, that exposed persons have the medical opinion necessary to occupy the post;

21) to ensure the information of the professionally exposed persons and the persons under preparation in accordance with the legislation in force;

22) to allow the representatives of the regulatory body in the field of nuclear and radiological activities - the Agency and the Supervision State Service of Public Health, empowered to exercise their right of control in accordance with the applicable provisions of Law no. 132 of 8 June 2012 and Law no. 10 of February 3, 2009 to ensure the participation of the radioprotection and radiological security officer or the expert attested to the radioprotection;

23) to fulfill any other obligations, that are necessary for the safe conduct of nuclear medicine.

90. Whenever there is a failure to comply with the included conditions in the radiological authorization, the holder of the radiological authorization shall:

1) investigate the causes, circumstances and consequences of the event;

2) take the necessary actions to eliminate the causes and limit the consequences, to notify the Agency and the Supervision State Service of Public Health in written form and within 24 hours from the finding;

3) decide to apply the necessary preventive and corrective measures, that have to be taken to reduce the probability of similar situations apparition;

4) stop work if it presents radiological or other risks that are unacceptable to the personnel or the patient;

5) follow the application of the measures, ordered by the Agency and the Supervision State Service of Public Health.

91. Responsibilities for radioprotection and radiological safety shall be clearly defined in writing, and the qualifications of each entrusted person with responsibilities shall be evidenced by the availability of appropriate diplomas, certificates and permits.

92. The radioprotection and radiological safety officer shall:

1) have a valid exercise permit for the field and specialty of the practice, in accordance with the applicable provisions of Law no. 132 of June 8, 2012;

2) be vested with sufficient authority to impose compliance with permit conditions and to stop work / activity if it is not carried out under radioprotection and safety conditions.

93. The radioprotection and radiological safety officer shall have the following duties and responsibilities:

1) to supervise, that the practice shall be performed and conditioned by compliance with the procedures and imposed conditions by the authorization and by the control reports;

2) to ensure that personnel are properly trained in the conduct of radioprotection and radiological safety of nuclear medicine;

3) to apply emergency procedures in the event of a radiological accident or incident;

4) to keep track of the metrological verification of the dosimetric control equipment and to verify how it is used;

5) organize the examination and training of all professionally exposed personal;

6) to prepare and revise periodically the laboratory's procedures for compliance with regulatory requirements;

7) to prepare and review work procedures so that exposure to radiation is kept as low as possible (ALARA);

8) to ensure that the instructions for use of the installations are known by the involved personnel;

9) to ensure the development and implementation of contingency plans in case of radiological emergency;

10) to ensure that radiation installations and sources are periodically checked, have a security certificate;

11) to identify and demarcate supervised areas and controlled areas and establish security and access control measures in these areas;

12) to ensure adequate individual dosimetric supervision and exposure to external and internal ionizing radiation, where appropriate, of professional exposures and maintain records;

13) to ensure radiological monitoring of the work environment and maintain records;

14) to investigate accidents / incidents, overexposure, and propose preventive and / or corrective measures;

15) to establish additional radiation protection measures for pregnant women;

16) to maintain up-to-date records of sources of radiation and radioactive waste by managing a register, including radioactive effluents, discharged into special sewerage;

17) to perform periodic radiological and radiological safety audits and maintain records of the results and / or corrective or preventive measures;

18) to identify the situations in which a radioprotection expert should be consulted;

19) to check, how to solve the noted problems by the certified radioprotection expert;

20) periodically to inform the holder of the radiological authorization of the work undertaken and the necessary corrective or preventive measures;

21) immediately to inform the holder of the radiological authorization of any event, accident / incident occurring in the practice of nuclear medicine.

94. Where more than one radioprotection and radiological safety officer is designated, the duties of each person shall be clearly defined in writing in an internal regulation, coordinated with the Agency and shall not overlap.

95. The radioprotection and radiological safety officer shall be responsible for the content, management and archiving of the technical documentation of the controlled area for which he has been designated.

96. The Nuclear Medicine Specialist / practitioner has the following responsibilities:

1) to provide patient radioprotection;

2) to prescribe and justify medical exposures in writing, taking into account relevant information from previous examinations;

3) to ensure that patient exposure is minimal to achieve the expected outcome, taking into account the baseline for medical exposures in the diagnosis;

4) to advise and clinically evaluate the patient;

5) to establish optimized protocols for diagnostic procedures, together with the medical physicist;

6) to assess any medical incident / crash and establish the notification procedure of the Public Health Supervision Service of the Ministry of Health, the Agency and other competent authorities;

7) to provide criteria for: examining pregnant or breastfeeding women and children, forensic medical procedures, medical examinations of occupational exposures, and biomedical investigations.

97. The medical officer has the following responsibilities:

1) to participate in the continuous assessment of used endowments in nuclear medicine practice and procedures;

2) to participate in the acceptance tests of the equipment;

3) to prepare the performance specifications of the equipment, taking into account the radioprotection requirements;

4) to supervise the maintenance of the equipment;

5) to design, implement and supervise compliance with quality of control procedures;

6) to perform the required calculations to estimate the absorbed doses in the relevant organs and the effective dose for the activity to be administered to the patient in accordance with the clinical protocol and / or international recommendations on the patient's internal dosimetry;

7) to participate in periodic quality control tests on dose calibrators, radiation monitors, gamma cameras and other relevant equipment in accordance with established procedures;

8) to calculate the work, that is going to be administered to children for a particular procedure;

9) to check, how the activity is measured prior to administration;

10) to participate in the investigation and evaluation of incidents and accidents;

11) to  contribute to the personnel training program.

98. The medical assistant has the following responsibilities:

1) to comply with radiation protection rules, procedures and operating instructions;

2) to ensure that patient identification is correct and that patient information is correctly recorded;

3) to provide information of patient on the procedure;

4) to provide information to persons , that are accompanying patients and personnel, that take care of  patients, who followed a nuclear medicine procedure;

5) to check and notify if some patients are pregnant;

6) to ensure that breastfeeding patients have been given information on the need to interrupt breast-feeding;

7) to ensure that the methods prescribed by the Nuclear Medicine Specialist are used to block the absorption of radiopharmaceuticals in organs not under investigation or methods to accelerate their excretion;

8) to make image acquisition and processing as appropriate;

9) to inform the radioprotection and radiological safety officer about incidents / accidents or any malfunctions in his / her activity;

10) to inform the radioprotection and radiological safety officer of found defects at installations and not use defective or inadequately maintained installations;

11) to inform the radioprotection and radiological safety officer, Nuclear Medicine Specialist, of any mismanagement;

12) to use correctly the dosimetric supervision means with which it is equipped;

13) properly to use individual and collective radiation protection means;

14) to report to the administration or management of the medical institution the cases of overexploitation of patients or staff according to the legislation in force;

15) to participate, according to his / her duties, in the established actions by the intervention procedure in emergency situations;

16) to participate in the training of persons in training in the practice of nuclear medicine.

The assistants may also have other responsibilities, which may also be the radiochemist, as appropriate:

a) preparation of radiopharmaceuticals according to the manufacturer's instructions and the applicable procedures;

b) performance of quality control procedures for the preparation of radiopharmaceuticals.

99. The persons in course of training should only perform their work under the supervision and guidance of the radioprotection and radiological safety officer or the nuclear medicine specialist.

100. The persons in course of training have the following duties and responsibilities:

1) to work only under the supervision of the radioprotection and radiological security officer or specialist nuclear medicine physician;

2) to comply with the radioprotection rules;

3) properly to use individual dosimetry, work and radioprotection equipment;

4) not to use defective or incorrectly maintained installations;

5) immediately to inform the radioprotection and radiological safety officer or Nuclear Medicine Specialist of any failure or nonconformity found.

CHAPTER VIII

THE OPERATIONAL RADIOPROTECTION SYSTEM

101. The holder of the radiological authorization is required to establish and maintain an operational radiological protection system.

102. The operational radiological protection system represents all the actions, procedures and instructions necessary to ensure the conduct of nuclear medicine practice under radioprotection and radiological safety.

103. The objectives of the operational radiological protection system are:

1) defining the responsibility of the radiological authorization holder for radioprotection, by adopting the organizational structures and the necessary procedures;

2) minimizing the risk of exposure to ionizing radiation of occupational exposures, other persons and persons subject to radiation diagnosis or treatment with open radioactive sources;

3) compliance with the ALARA principle;

4) establishing quality assurance requirements in operation, including verification of radioactive sources, radiological installations and dosimetric control equipment;

5) establishing measures for the physical protection of radioactive sources;

6) compliance with the requirements and limits set out in the radiological authorization.

104. The operational radiological protection system involves the following measures:

1) observance of the general principles of radiation protection, foreseen in the radiation protection regulations in force, elaboration and implementation of the procedures that ensure the safe conduct of the nuclear medicine practice;

2) practice only by the persons holding the valid authorizations and permits for the respective activity;

3) Written assignment of radioprotection and radiological safety responsibilities for each exposed person.

105. The operational radiological protection system is based on a general practice of practice which must establish:

1) organization and forms of practice;

2) controlled and supervised areas;

3) the obligations and responsibilities of all the involved factors;

4) administrative relations between the involved factors;

5) system documents: keeping, archiving, broadcast, withdrawal;

6) necessary records.

106. The operational radiological protection system shall be supplemented by the following operational procedures and instructions:

1) individual dosimetric control procedure;

2) the radiological monitoring procedure of the workplace;

3) clinical dosimetry procedure;

4) diagnostic procedures;

5) the procedure for the recording, movement and storage of radiological installations and radioactive sources;

6) quality control procedures for each of the used equipment in the nuclear medicine laboratory;

7) the procedure for controlling access to the controlled area;

8) procedure for the collection and elimination of radioactive waste;

9) the procedure for the use of installations and equipment;

10) the procedure for investigating and reporting non-normal overexposure and exposure;

11) the radiological emergency intervention procedure.

107. The control of radioprotection insurance in nuclear medicine subdivisions is performed by the radioprotection service and / or the person responsible for radiation protection, appointed by order by the administration of the medical institution. States and technical insurance are set by the administration of the medical institution, depending on the volume and activity of the radiopharmaceutical preparations.

108. The regulation of the radioprotection service is approved by the administration of the medical institution and coordinated with the Agency and the State Health Supervision Service of the Ministry of Health.

109. The composition, the volume, the periodicity and the specific places of radiological control are established by the administration of the medical institution, coordinated with the Agency in the authorization process and by the State Health Supervision Service of the Ministry of Health.

110. Radiological control is planned, in cases of modification of the technological process, the presence of data on violation of radiopharmaceutical preparations and in cases of accident.

111. When experimenting with radiopharmaceuticals, the implementation of new methods of radiological control is performed in the volume and periodicity necessary to determine the optimal regime of routine radiological control.

112. Radiological control uses dosimetric and radiometric means for verified metrological gamma and beta radiation indicated for such measurements with an error of not more than ± 20%.

113. The results of radiological control shall be recorded in registers including workplace diagrams showing the location of radiopharmaceutical preparations and radioactive contamination control points.

114. Procedures in the operational radiological protection system that refer to the improvement of the quality and results of the investigations of patients will be periodically checked by clinical audits, which will be carried out in accordance with the specific regulations of the Ministry of Health and recognized by the Agency.

CHAPTER IX

RADIOPROTECTION INSURANCE

OF PROFESSIONALLY EXPONENTS

115. The general principles for the radiological protection of professionally exponents in subdivisions are a system of requirements for personnel, rooms, technologies, and methods and means of radioprotection.

116. The radioprotection of professionally exponents in subdivisions is ensured by:

1) limitation of activity in the field of ionizing irradiation sources according to age, gender, health status, previous levels of occupational exposure and / or accidental irradiation;

2) sufficiency of collective radipprotection equipment and limitation of working time with ionizing radiation sources;

3) strict observance of the requirements and recommendations for personnel, which are described in the job descriptions and the normative-technical documentation;

4) use of individual radiation protection means;

5) organizing the radiological control of supervised and controlled areas;

6) effective measures to prevent and eliminate radiological accidents.

117. In order to reduce the level of internal exposure of personnel, it is necessary:

1) to exclude access to premises with open radioactive sources to people, who do not participate in the work with these sources, including patients, employees and others;

2) optimizing the distance between the source and the professionally exposed, including the distance between employees and patients, who are investigated with radiopharmaceuticals;

3) decrease of the time of exposure of the professionally exponents in the field of ionizing irradiation sources, including the duration of contact time with the administered patients with the radiopharmaceuticals, avoiding the violation of the medical procedures and diminishing the quality of the diagnosis;

4) reducing the activity of radioactive sources, in the field of which there are professionally exponents;

5) the use of stationary radiation protection equipment, lead mountings, protective shields, screens, containers and speakers;

6) use of tools for remote manipulation of sources and radioactive waste;

7) Mandatory administration of radiopharmaceuticals with disposable syringes.

118. In order to exclude the incorporation of radionuclides and to reduce the levels of internal exposure of professionally exponents, it is necessary:

1) performing all work in nuclear medicine (laboratories) with the use of personal protective equipment (gown, cap, gloves, lightweight non-woven footwear);

2) the use of individual means of protection in the cleaning of rooms, where activities with open sources of ionizing radiation, WCs for patients are performed;

3) the use of a set of individual means of protection and means of breathing protection to eliminate the consequences of accidents with ionizing radiation;

4) handing out special contaminated clothing for disabling special laundry and periodic exchange of special clothing not less than once every two weeks;

5) the use of individual disposable protection means;

6) to attract special attention to the non-admission of the spread of radioactive contamination outside the controlled area.

119. The activity with sources is allowed for persons, who have reached the age of 18, have training in the respective field, have passed the training and examination of the knowledge for radioprotection, possess the exercise permit specific to the performed activities, know the provisions of the normative acts and the instructions in force in the medical institution, according to the applicable provisions of Law no. 132 of June 8, 2012.

120. The medical institution's administration must perform mandatory medical control on hiring and periodically (once a year) the occupational exposure to ionizing radiation. In activity with radiopharmaceuticals are admitted for persons, who do not have medical contraindications. These requirements also apply to persons, who prepare staff for nuclear medicine departments (laboratories) in the course of specialization, training, residency.

121. When detecting deviations in the health status of personnel that do not allow radioactive activity, the administration of the medical institution shall investigate each individual case and decide to temporarily or permanently transfer to another job, excluding contact with radioactive sources.

122. Investigations with radiopharmaceutical preparations, accompanied by complicated manipulations, which do not fall within the functional duties, may also be attended by specialists, who will be included in the category of professionally exponents.

123. The staff is obliged to know and comply with the requirements of these Regulations, the rules of radioprotection, labor hygiene, fire safety and security. When detecting violations of the protection means, the personnel is obliged to notify the administration of the medical institution.

124. It is forbidden to perform works that are not provided for in the officially approved instructions on radioprotection and the work safety technique.

125. It is forbidden to perform quality control, installation, repair, adjustment of facilities during patient investigations.

126. Staff are required to have first-aid delivery methods, to have the addresses and telephones of institutions, persons to be contacted in case of accidents, possible incidents.

127. In the event of an accident or a radiological incident, the radiological service personnel shall act in accordance with the approved instructions for dealing with the consequences.

CHAPTER X

PERSONNEL TRAINING

128. The personnel pursuing the practice of nuclear medicine shall have:

1) general training appropriate to the job;

2) training in the field through appropriate nuclear medicine courses, according to the regulations of the Ministry of Health;

3) training in radioprotection, through approved courses by the Ministry of Health and recognized by the Agency and other competent bodies;

4) extensive knowledge of the operational radioprotection system;

5) Practical training in the practice of nuclear medicine with sources of ionizing radiation and procedures for intervention in case of radiological emergency;

6) Appropriate exercise permits to practice deployment.

129. The exploitation of nuclear medicine apparatus is the admitted persons, who have reached the age of 18 and are assigned as professionally exposed workers, have appropriate papers in the field, have been trained and evaluated their knowledge of normative acts and instructions in force, which give the right to work in the field of nuclear medicine.

130. The programs of internal training courses in organized radioprotection by the holder of radiological authorization shall be subject to the topic, approved by the Ministry of Health and recognized by the Agency.

131. The holder of the radiological approval is required to ensure appropriate training in the radiological safety of the professionally exposed personnel and to recycle it at least once every 5 years through an approved training system by the Ministry of Health and recognized by the Agency.

132. The holder of the radiological approval is required to instruct the persons in training process prior to their use of installations with ionizing radiation sources. Training data will be entered in a special register in accordance with Annex no. 3 to these Regulations. The instructions for radioprotection and radiological safety shall be attached to the Register. The register is kept permanently.

133. The holder of the radiological authorization shall ensure that the radioprotection and radiological safety officer and the occupational exposure to ionizing radiation, is known the following:

1) the conditions and limits of the radiological authorization and sanitary authorization for the operation of the objective;

2) operational radioprotection System, developed at the organization level;

3) instructions for the use and maintenance of installations and used equipment in nuclear medicine, installations and equipment for the measurement and detection of ionizing radiation;

4) the associated risks with the practice;

5) procedures and instructions for work;

6) procedures for intervention in case of radiological emergency.

134. The holder of the radiological authorization shall establish and maintain records of the training of involved personnel in the conduct of nuclear medicine. These records shall contain at least the following information:

1) the name and surname of the person attending the training;

2) the identification data of the institution that organized the course;

3) course duration;

4) course theme;

5) the copy of the certificate of graduation of the training program;

6) copy of the exercise permit for that practice.

CHAPTER XI

EXPOSURE OF PUBLIC PERSONS

135. After administration of the radiopharmaceutical preparation, the patient himself becomes a source of external irradiation and in some cases of internal irradiation of other people.

136. The security and radioprotection of persons, who are in episodic or regular contact with the patients, who have administered the radiopharmaceutical preparation is ensured:

1) thorough and strict observance of the given instructions and recommendations by the radiologist himself to the patient, his / her companions and relatives;

2) limiting as far as possible the maximum possible contact time and increasing the distance between the patient and other people, especially children;

3) temporary interruption of breastfeeding if the mother was given the radiopharmaceutical preparation.

137. The access of involved persons in supporting and caring for patients and their family members in controlled areas and supervised areas is controlled and monitored.

138. The involved patients in supporting and performing for patients and their family members must be accompanied in the controlled areas by a trained person with regard to the radiation protection and security measures in that area and be informed of the rules to be followed to respect them.

139. Individual dosimeters (digital, direct reading or area dosimetry) will be used to control the dose of individuals, participating in the support and care of patients as well as their family members and the received dose will be recorded at each visit and for each of mentioned persons.

CHAPTER XII

INDIVIDUAL MONITORING AND EXPOSURE EVALUATION

140. The holder of the radiological authorization is required to ensure that occupational exposure estimation of workers is performed on the basis of systematic individual dosimetry monitoring of all professionally exposed workers to ionizing radiation and to organize their medical supervision for the assessment of the permanent compatibility of the exposed worker to the conditions of work.

141. The Individual Dosimetry Service issues individual dosimeters for use and the radiological authorization holder submits them for evaluation within the established deadlines.

142. The individual dosimetric monitoring of exposed workers to occupational exposure will demonstrate compliance with the basic described radioprotection principles in current legislation and these Regulations.

143. The State Public Health Service and the Agency are entitled to request the holder of the radiological authorization to ensure individual dosimetric monitoring and other persons, that are not included in the category of professionally exposed workers but working with sources of ionizing radiation.

144. The system and type of individual dosimetric monitoring of exposed workers, the required procedures, method, periodicity, type of dosimeters shall be established and approved by the Agency jointly with the Ministry of Health in the authorization process and shall be performed according to the requirements of the radioprotection for the calculation of the effective dose (E, [Sv]) or the equivalent dose to the skin or extremities (HT, [Sv]).

145. The operational quantity for individual dosimetry monitoring for radiopharmaceutical investigations is the equivalent of the personal dose (Hp (d), [Sv]).

146. Periodicity of individual dosimetric monitoring is set from 1 to 3 months depending on the performed activity, the practice-related prediction of the exposure, the characteristics of the individual dosimeter and the minimum exposure detection level of the dosimetric system.

147. The individual dosimetry monitoring is performed through the individual dosimetry service, that is accredited in the National Accreditation System and the National Health Assessment and Accreditation Council within the Ministry of Health and recognized by the Agency. The State Health Supervision Service will analyze quarterly, half-yearly and yearly the indices of performing individual dosimetric monitoring.

148. Individual dosages resulting from external exposure will be determined by using devices (individual dosimeters) of individual dosimetry systems (thermoluminescent, photoluminescent, doubling in required cases with integral electronic or direct reading dosimeters).

149. The individual issued dosimeter to the professionally exposed worker shall be performed only by the person to whom it was handed.

150. In case of loss or deterioration of the individual dosimetry, the individual dosimetry service will perform the assessment and registration of the dose for the respective worker by alternative methods (based on the employee's work-related calculations under similar working conditions, or by assessing the received dose levels in earlier periods). The loss (deterioration) of the individual dosimeter and dose estimation are reported to the Public Health Supervision Service of the Ministry of Health and the Agency.

151. The holder of the radiological authorization shall disclose all individual dosimeters within the deadlines set for processing. The failure to comply with the terms of presentation of individual dosimeters qualifies as a disciplinary offense. The accredited individual dosimetry services have to provide additionally the radiological authorization holder with a dosimeter, in exchange for those presented to read the accumulated parameters.

152. In the case of activity in non-homogeneous fields, when the individual protective apron is used, the accredited expert in radioprotection may decide to carry an additional dosimeter above the "tandem" apron with the used one. In case of emergency or accidental exposure is performed the individual dose monitoring or evaluation, taking into account each case.

153. The individual doses of the exposed worker are confidential. Access to dose information is provided by the person being monitored, the Radiological Authorizing Officer / Radiation Protection Officer, the Public Health Supervision Agency and the State Health Supervision Service, the Occupational Pathology and Occupational Medicine Medical Services to assess the implications for human health. The accredited individual dosimetry service and the radiological authorization holder are responsible for the veracity and confidentiality of the data. The data are released in the deadlines and forms established by the Ministry of Health.

154. In case of transfer (dismissal) of the exposed worker, the holder of the radiological authorization shall address to the individual dosimetry service for the issue of the extract, indicating the data on the received doses during the activity period.

155. For each worker who is professionally exposed, the results of the individual monitoring are recorded, with their registration in the individual dossier records and the national register of the exposed workers, and their duplication in digital form. This record shall be kept centrally at the individual dosimetry service of the Public Health Supervision Service, and at the local level - at the radiological authorization holder. The retention period is up to 75 years of age and not less than 30 years from the time of cessation of activity in the field of ionizing radiation.

156. The individual dosimetry service shall present, within the set deadlines, the generalized data and relevant information of the individual dosimetric monitoring with the actual doses to the State Public Health Service and the Agency. The same data shall be submitted to the professional pathology sections at their request so that they can perform medical exams prior to hiring or qualification of professional exposures, as well as supervising subsequent exposures after employment.

157. The holder of the radiological approval shall establish and maintain a procedure for the investigation and reporting of overexposure and non-normal exposure. The procedure must provide for:

1) prompt notification of the affected person;

2) the immediate announcement of the Agency and the State Public Health Supervision Service, and of the external worker's employer, if the concerned person is an external worker, on overestimated overexposure and preliminarily assessed doses;

3) to draw up a report on the event within 10 days of the commencement of the investigation and to submit it to the Agency and the Public Health Supervision Service;

4) investigating the case of overexposure by the Public Health Supervision Service and the Agency.

158. The report shall contain the following information:

1) the results of the investigation on the basis of which the preliminary value of the received doses was established;

2) the circumstances in which overexposure was occurred;

3) evaluating the received dose, based on all available data, including on the basis of the results of individual dosimetric measurements;

4) taken measures to prevent such overexploitation.

159. Occupationally exposed persons, who have been overexposed over the specific control levels of the workplace are entitled to continue to work in controlled areas if there are no medical contraindications.

CHAPTER XIII

RADIOLOGICAL MONITORING OF AREA

AND THE WORKING ENVIRONMENT

160. The holder of radiological authorization shall ensure that the radiological monitoring of the workplace of the exposed workers to the workplace is performed. Radiological monitoring of workplaces for controlled areas and adjacent areas of controlled areas is performed by daily measurement of dose rates due to external exposure, with indication of radiation level. The monitoring of working conditions of workers, that are exposed to occupational exposure should include the individual doses of exposure to ionizing radiation.

161. The workplace monitoring of exposed workers to occupations in particular practices may be supplemented with dose measurements by devices (dosimeters) of the used type for individual dosimetry monitoring. The devices will be placed at points with the maximum dose rate, or at the most frequented places in controlled or supervised areas.

162. In order to assess the radioprotection status in the institutions, that perform radiopharmaceutical investigations, general dosimetry control is mandatory.

163. All used instruments, devices, dosimeters to monitor workplaces must be metrological checked. The operating status of the warning devices will be checked daily.

164. The radiological authorization holder shall record the results of measurements of the radiation fields in the controlled and supervised area, performed at the maximum exposure points set in a work area monitoring plan, approved by the Agency and the State Health Supervision Service .

165. For beta and beta-gamma radionuclides, the results of the contamination measurements are compared with the admissible levels of contamination of work surfaces, skin, special clothing and individual protection means, according to the applicable radioprotection regulations.

166. For gamma radionuclides, control of the level of contamination is performed by measuring the flow rate of the absorbed dose into the air at a distance of 10 cm from the contaminated surface.

167. Measurements of the absorbed dose rate into the air at workplaces of professional exposures shall be performed at three levels: 0,1, 0,9 and 0,5 m from the floor surface to the maximum possible activity of the source for the given procedure. Three measurements are performed at each point with results mediation.

168. The dosimetric control of the radioactive waste shall be performed upon their entry into the storage repository, their settlement and transmission to the centralized storage. The dosimetric and radiometric control of special clothing and household articles are carried out prior to handing over to specialized laundries for the purpose of sorting and separating objects to be subjected to radioactive disintegration.

169. The results of the radiological control are compared with the dose limit values ​​or the level of surface radioactive contamination and the control levels according to the radioprotection regulations in force. Exceeding the control levels is analyzed by the radioprotection service, which informs the administration of the medical institution. If the dose limits set for professional exposure are exceeded, the healthcare institution's administration announces the Public Health Supervision Service of the Ministry of Health and the Agency.

170. It is forbidden to carry out dosimetric measurements at workplaces of professionally exponents without the use of individual radiation protection devices.

171. The measurements of dose rate measurements at workplaces serve as a basis for establishing working arrangements for professionally exponents during routine work.

172. The planned dosimetric dosing control in workplaces is performed no more than once a year.

173. The radiological monitoring of the working environment shall include measurement of dose rates due to external exposure and contamination with radionuclides.

174. The dose rate monitoring due to external exposure is mandatory for all controlled areas.

175. The monitoring of contamination is mandatory for:

1) all work surfaces, including their interior, all tools, equipment, floor and other removed objects from the area;

2) the personal and protective clothing, including footwear, especially when the person is in a controlled area. For this purpose, a radiometer for checking superficial contamination of the body (hands) will be placed at the exit of the area;

3) the clothing and linen of suspected patients of contamination;

4) the hands of people using open radiation sources will be monitored by placed instruments near the wash area. Monitoring will extend to other areas of the body if suspected contamination.

176. The radiological monitoring of the area and the working environment shall be performed either by suitably trained personnel with the provided equipment or by a qualified external entity.

177. Measurement results will be recorded during practice and will be used, if necessary, to estimate individual doses. The minimum retention time of the intended registrations is 10 years.

178. The holder of the radiological approval shall establish and maintain a procedure for the radiological monitoring of the working environment. The procedure should contain the following information:

1) the measuring equipment and methods, including methods for checking the dosimetry control apparatus prior to use;

2) the Nuclear Medicine Laboratory Plan and the measuring points;

3) the measurement procedure;

4) responsibilities;

5) record and interpret the results;

6) the corrective measures to be taken if the established levels of investigation are exceeded.

179. The records shall contain the following information:

1) the sketch of the laboratory, including the indication in the plan of the adjacent areas of the treatment rooms, the investigation or storage rooms of the radioactive sources and the points where the measurements are made;

2) the date on which the measurements were made;

3) the name of the person who performed the measurements;

4) the type of used dosimetric or radiometric control apparatus and its identification data: the serial number, the year of manufacture, the date of the last metrological verification;

5) the ​​obtained values for the dose rate and / or radioactive contamination.

180. The radiological authorizing holder shall be provided with at least one radiodetectometer and a contaminator or with a single dosimetric control device, if it can work in a radionuclide meter and a contaminator.

181. The used dosimetry control equipment shall be appropriate to the types of radiation, emitted by used radioactive sources in the conduct of nuclear medicine practice.

182. The radiometer shall be calibrated in dose rate units, capable of indicating at least values between 0.1 mSv / h (microSievert / hour) and 1.0 mSv / h (milli-sieve / hr) accuracy of at least +/- 30% and have a sound signal that can be reduced when working in close proximity to patients.     The radiometer to detect surface contamination is to be gamma and beta type.

183. The radiometer for surface contamination detection shall be capable of detecting gamma and beta radiation.

184. The dosimetric control equipment shall be checked at the beginning of each working shift.      This verification shall include:

1) checking the condition of the battery packs;

2) measurement of the natural radiation background;

3) radiometric or dosimetric equipment response when approaching a container with a gamma radioactive source or using a test radioactive source.

185. The verification performing is the responsibility of the radioprotection and radiological security officer or another person designated in writing by the radiological and radiation safety officer.

CHAPTER XIV

HEALTH SUPERVISION STATE OF

PROFESSIONALL EXPOSED WORKERS

186. The supervision of the annual medical examinations of professionally exposed workers to ionizing radiation shall be ensured in accordance with the general principles governing occupational medicine by specialists of the Public Health Supervision Service of the Ministry of Health. The medical supervision of the professionally exposed workers comes back to the clinicians in the medical institutions with competence in the field of professional pathologies. It is intended to enable the health of supervised workers to be established and their ability to carry out their work.

187. The competent physician and the occupational pathology subdivisions of the medical institutions, established by the Ministry of Health shall have access to any relevant information they require including information on environmental conditions at the workplace. For this scope, at the request of the medical institution, public health centers perform the supervision of the environmental conditions at the workplace.

188. The holder of the radiological authorization shall organize and ensure the medical supervision of the exposed workers to ionizing radiation in accordance with the established regulations by the Ministry of Health. The radiological authorization holder shall draw up a protection and prevention plan which includes technical, sanitary, organizational and of a different nature, based on professional risk assessment.

189. The health supervision programs of professionally exposed workers are based on the general principles of occupational hygiene and are intended for primary and permanent assessment of the health status of employees for the performance of their service obligations in accordance with the legislation in force.

190. The medical supervision shall include a compulsory medical examination of the worker , performed prior to employment in the field of ionizing radiation and the classification of workers who are professionally exposed and a periodical examination once a year.

191. After the approval by the State Public Health Supervision Service of the lists of exposed workers and the medical examinations on recruitment and periodical, the commission, in the composition of the representative of the professional pathology service, the Territorial Public Health Center, the trade union organization (if exists) the representative of the holder of the radiological authorization, signs the final act of the medical examination results.

192. The conclusions regarding the professional aptitude (apt, apt with certain conditions, inappropriate) after the medical examination are recorded in the medical file in the approved form by the Ministry of Health. By transferring the person to another activity or enterprise, after the termination of the activity, the medical record with the necessary additions shall be sent to the respective medical institution.

193. It is forbidden to work with sources of ionizing radiation of employees, who have not performed the preventive and periodic medical examinations and do not have a permit for activity under harmful conditions.

194. In case of detecting deviations of the health indicators of the personnel, which do not allow to continue its activity, the administration of the medical institution will study each individual case and will take the decision to transfer temporarily or permanently to a job that would avoid exposure to ionizing radiation. In case of the "inappropriate" result, the person will be removed from the environment with ionizing radiation.

195. In the case of accidental exposure at high doses (of the order of 0.2-0.5 Sv (Sievert) or higher), it is necessary to carry out the repeated special medical examination, the application of the decontamination measures and the curative emergency treatments in relation to the associated risk irradiation.

196. The holder of the radiological authorization shall retain the results of the medical supervision of exposed workers to ionizing radiation, reports and records after passing the levels, overexposure and accidental exposure.

197. For each worker of Category A, an outpatient medical file shall be opened which shall be updated as long as the worker remains in that category. After the cessation of activity, the medical record is kept until that person reaches or has reached the age of 75 but in no case less than 30 years after the end of the activity, involving exposure to ionizing radiation.

198. The outpatient medical record must include information on the nature of the post, the results of the medical check-up before recruitment or classification in Category A, regular health reviews.

CHAPTER XV

REQUIREMENTS FOR RADIOACTIVE SOURCES

199. The open administered radioactive sources in nuclear medicine for diagnostic purposes are radiopharmaceutical products.

200. In vitro diagnostic RIA kits (set of specific in vitro diagnostic markers) are not considered to be radiopharmaceuticals. These are specific pharmaceutical markers for the investigated organs, allowing them to be explored for diagnostic purposes.

201. The introduction of pharmaceutical products containing radioactive materials and which may be used in nuclear medicine and the quality requirements, that they are going to meet are set by the Ministry of Health, which authorizes their use for diagnosis.

202. Pharmaceutical products and RIA kits, containing radioactive materials require the radiological security certificate and health certificate, prior to their use.

203. The radiopharmaceutical products and RIA kits, containing radioactive materials shall meet the applicable radiation and pharmaceutical standards for:

1) radionuclide purity;

2) specific activity;

3) radiochemical purity;

4) chemical purity;

5) pharmaceutical properties: toxicity, sterility, pyrogenicity.

204. The holder of the radiological authorization will establish a quality control procedure for pharmaceutical products and RIA kits, containing radioactive materials for their preparation for administration. The procedure will include:

1) written instructions for the reception of pharmaceutical products and kits containing radioactive materials or generators: verification of the correspondence between the documents on the packaging and the written ones, visual inspection, activity measurements, appropriate storage;

2) written instructions for the preparation and administration of pharmaceutical products and RIA kits, containing radioactive materials as specified by the manufacturer;

3) quality control tests on radionuclide generators, containing: the monitoring of the daily activity of the eluate as an indicator of generator operation, the measurement of the passage of the parental radionuclide in the eluate, chromatographic tests for the percentage of radiotracer marking;

4) monitoring the preparation conditions: hygiene, contamination, aseptic;

5) way of labeling and marking the bottles: the type of radionuclide, compound, the activity at that date, the volume of the solution;

6) used equipment: radiochromatography equipment, sterilization control equipment, dose calibrators;

7) records of patient activity: date of injection / administration, patient name, radiopharmaceutical, injectable / administered activity.

205. The pharmaceutical products and RIA kits, that are containing radioactive materials are going to be labeled appropriately. The label shall contain the following information:

1) the radionuclide and the chemical name of the preparation;

2) the total activity / date at which it was measured;

3) the expiration date;

4) the manufacturing batch number;

5) the volume of the solution in the case of liquid products;

6) the name and address of the manufacturer.

CHAPTER XVI

REQUIREMENTS FOR EQUIPMENT USED IN NUCLEAR MEDICINE

206. The used equipment in nuclear medicine shall be provided with the radiological security certificates, issued by the Agency.

207. The holder of a radiological authorization shall establish and maintain a program of verifications and maintenance of the equipment to take into account the recommendations of the manufacturer and the supplier.

208. The program of verifications and set up maintenance work shall contain:

1) the acceptance / reference verifications that are performed before commissioning to ensure that they operate safely and as specified by the manufacturer.

2) the regularity of the regular checks and / or calibrations of the provided equipment in the established manner and in accordance with the legislation in force;

3) the routine checks (daily, weekly or monthly) that are executed by designated persons in writing by the radioprotection and radiological security officer at the beginning of the work program.

209. Any failure detected in the operation of the equipment shall be notified to the radioprotection and radiological safety officer.

210. The results of routine and periodic verification shall demonstrate the compliance with the results of the acceptance / reference verifications. The results of the acceptance / reference verifications and routine verifications, including information on defects found and performed repairs for each type of equipment, are going to be kept in a special register with the holder of the radiological authorization.

211. After each repair or replacement of a component that could affect the specified parameters by the manufacturer, the acceptance verifications shall be re-established and the notified Agency.

212. The radiological authorization holder shall develop and implement the quality control procedures.

213. The quality control procedures shall include, for each type of used equipment in the nuclear medicine laboratory, at least the following: the type of required verifications for each type of equipment, the specific test procedures, the frequency of the measurements, the data processing, the interpretation of the results, the acceptability limits , means of measurement, radioactive sources of test and other materials necessary to perform those verifications, the persons responsible for carrying out the quality control tests. The recommended minimum checks for dose calibrators, counting systems, gamma cameras, SPECT (Single-photon emission computed tomography) and PET systems are given in Annex 4 to these Regulations.

CHAPTER XVII

PHYSICAL PROTECTION, MANAGEMENT

AND EVALUATION OF RADIOACTIVE SOURCES

214. The holder of a radiological authorization shall establish and maintain a procedure for the management, recording, movement and storage of radioactive sources. The procedure should contain the following information:

1) the general and operative management and accounting system;

2) the necessary measures to prevent the theft, loss, destruction and use by unauthorized persons;

3) the way and rules for removing radioactive sources from the warehouse, entering the warehouse, tracking the place of use;

4) the recordings of the movement and consumption of the radioactive sources;

5) the responsibilities;

6) the sanctions.

215. The radiological authorizing holder shall establish and maintain a primary and operative record of all radioactive sources that he uses in the course of practice to result in the consumption of radioactive sources.

216. The evidence of the use of radioactive sources shall contain the following information:

1) radiopharmaceutical name / radionuclide designation;

2) activity on production date;

3) manufacturer and supplier of the radioactive source;

4) date of entry into account;

5) series and batch number of manufacture;

6) the place of storage;

7) the date and time of release from storage or consumption;

8) the names of the persons who are designated to use them;

9) date and time of re-entry into storage, if it’s appropriate.

217. The evidence of consumption of radiopharmaceuticals and generators shall be made according to the indicated registers in Annexes no. 5, no. 6, no. 7 and no. 8 to these Regulations. For each type of radiopharmaceuticals in vivo, they set different pages. The pages of the registry must be numbered, sewn and stamped. The records are kept permanently.

CHAPTER XVIII

CLASSIFICATION OF ZONES

218. The classification of the zones shall be made by an appropriate assessment of the associated radiological risk with the use of ionizing radiation sources, taking into account possible routes of exposure to ionizing radiation.

219. The preparation, storage, temporary storage and radiopharmaceutical injection rooms, imaging cameras, waiting rooms for patients receiving radiopharmaceuticals for diagnostic purposes, salons for treated patients, radioactive waste dumps are going to be classified as controlled areas.

220. For the classification of the areas, the criteria according to Annex no. 9 to these Regulations.

221. The control area requirements:

1) warning by signs of radioactive danger;

2) existence of appropriate instructions at the entry points in the controlled areas;

3) existence of procedures specific to each controlled area;

4) access control, through:

a) the administrative procedures;

b) the physical barriers with communication systems, if it’s appropriate;

c) the light and sound warning devices.

5) availability of appropriate means for changing clothes, monitoring of exposure and contamination, placed staff decontamination at the entrance / exit of the controlled area.

222. The area around the controlled area is considered the supervised area. The supervised area should be described in the plan. The requirements for supervised areas:

1) warning signs;

2) control access of people in the population.

223. In order to demonstrate the fulfillment of the requirements for the classification of the controlled and supervised areas, the holder of the radio-logic authorization has the obligation to carry out the strict delimitation of the respective areas.

224. The fitted warning signs at all access points shall clearly indicate the delimitation of controlled and supervised areas.

225. The access and the carrying out of works in the controlled area are allowed:

1) the professionally exposed persons to the practice of nuclear medicine;

2) the patients, who have been prescribed the procedure with radiopharmaceutical preparations;

3) the persons, participating in the support and care of patients, who have received radiopharmaceuticals for diagnosis.

226. The Radiological Authorization Holder is going to provide to persons having access to controlled areas: the appropriate means for changing clothes, adequate radioprotection equipment, monitoring of ionizing radiation exposure and contamination and, if it’s appropriate, decontamination, placed at the entrance / exit of the controlled area.

CHAPTER XIX

REQUIREMENTS FOR THE STORAGE

OF RADIOACTIVE SOURCES

227. The storage of radioactive sources is to be carried out only in a specially arranged room where other products or tools necessary for carrying out the practice are not stored, or in a storage container, protected from the risk of flooding and which can ensure them against degradation mechanical and / or due to the action of environmental factors, being provided with:

1) adequate insurance against theft;

2) radiodebitometer in the radioactive source or radiation with alarm threshold;

3) appropriate screens against gamma radiation, so that the dose rate for professionally exposed personnel, the only one having access to the deposit, does not exceed 7.5 μSv / h;

4) appropriate warning systems and, if it’s appropriate, means of extinguishing the fire;

5) adequate ventilation systems in accordance with the provisions of these Regulations.

228. The access door to the radioactive source or storage container shall be clearly marked with the "Radiation hazard" symbol and fitted with an unauthorized opening alarm system, connected to the equipped central alarm panel with a telephone connection with the maintenance representatives of public order.

229. In the case of radioactive waste storage, a sufficiently large special compartment for the safe storage of solid radioactive waste, exhausted generators and containers shall be provided if there is no specially arranged room for the storage of radioactive waste.

230. The access of authorized persons to the radioactive waste deposit shall be made only after the registration in the special register, indicating the date, the name and surname of the person and the type of removed container or placed in the deposit. The keys from the radioactive source storage shall be kept and released by the person responsible for the physical protection of designated radioactive material by an administrative act.

CHAPTER XX

REQUIREMENTS FOR RADIOPROTECTION MEANS

231. The laboratories and areas, where radioactive sources are used shall be provided, where appropriate, with at least the following:

1) radioprotection devices:

a) devices, increasing the distance between the professionally exposed person and the radioactive source: forceps, pliers, manipulators;

b) shielded containers for temporary storage of open sources of radiation during work;

c) double wall containers with an unusable exterior wall for liquid radioactive samples;

d) trays for working with open sources of radiation;

e) automatic or disposable pipettes or syringes;

f) screens for syringes and ampoules;

g) walls or castles of lead;

h) lead glass screens;

i)  perspex screens for beta radiation shielding.

2) personal protective equipment:

a) laboratory gowns and footwear, waterproof gloves, masks, skins;

b) radioprotection: lead aprons, lead gloves, waterproof gloves thick enough to work with beta radionuclides, respiratory protection masks in case of radiological emergencies.

232. The individual radioprotection equipment shall be used in areas where there is a risk of contamination both to protect the wearer's body or clothing and to prevent the spread of radioactive contamination in other areas.

233. At all individual protection and radioprotection equipment, contamination shall be checked and, where appropriate, removed before leaving the controlled area.

CHAPTER XXI

DIAGNOSTIC PROCEDURES. BREASFEEDING.

234. The holder of the radiological authorization shall establish and maintain diagnostic procedures to ensure the following:

1) the correct identification of patients;

2) the written protocols for each type of investigation, so that the obtained information is maximum, taking into account the reference level for that investigation, indicated in Annex 10 to these Regulations;

3) the accurate recording of details for the patient;

4) the choosing the right purchase conditions so that image quality is optimal: collimator selection, collimator contamination, energy window, acquisition matrix size, acquisition time, collimator tilt, SPECT parameters, PET, magnification factor;

5) the use of methods to block the absorption of the radiopharmaceutical in organs not under investigation and, respectively, to accelerate excretion, if it’s appropriate;

6) the avoidance of increased radioactive concentration in the workplace by external elimination or through collection on filters or in a leak-free container in cases where radioactive gases such as 133 Xe or radioactive aerosols are administered;

7) the informing, in writing and orally, the persons to whom radiopharmaceuticals have been administered and their attendants about the necessary precautions and the involved risk in an activity exceeding 300 MBq of 131I or 99Tc.

235. The holders of the radiological authorization shall:

1) ensure that reference levels for medical exposures are established, respected and do not exceed the admitted values and that they are used by involved healthcare professionals in the practice;

2) provide corrective measures if the doses or activities of administered radiopharmaceuticals do not provide good diagnostic information so that they do not result in a benefit for the patient;

3) have a revision of the procedures, if the used activities exceed the reference levels in order to ensure optimal patient protection and maintain adequate levels of good practice.

236. The radiologist is going to check whether the patient subject to radiopharmaceutical procedure is breastfeeding before any nuclear medicine procedure commences. In the affirmative, special consideration should be given to justifying and optimizing medical exposure, according to radiation protection rules, taking into account both the exposure of the mother and the child.

237. The radiologist shall provide the patient with written instructions on the interruption of breastfeeding in the required cases.

238. These indications ensure that the effective dose of the fetus or child being breastfed, as a result of maternal exposure is less than 1 mSv.

239. The table in annex no. 11 of these Regulations shall be used to establish periods for the cessation of breastfeeding. For other radiopharmaceuticals other than those in this table will be developed and approved the special written procedures.

CHAPTER XXII

CLINICAL DOZIMETRY. DOSING CONSTRUCTIONS

240. The holder of the radiological authorization shall institute and maintain a clinical dosimetry procedure to ensure that:

1) physical parameters and quality of radiopharmaceuticals, prepared for diagnostic use are checked;

2) the activity to be administered is determined, measured and recorded in the special register in the moment of administration for diagnostic procedures;

3) the absorbed doses in the relevant organs and the effective dose are determined and documented for the activity to be administered according to the clinical protocol.

241. The holder of the radiological authorization shall apply constraints to any received dose by the involved persons in supporting and caring for patients and their family members.

242. The constraints applied will not lead to higher doses than the following:

1) 1 mSv for children and for embryo / fetus;

2) 3 mSv for adults up to the age of 60;

3) 15 mSv for adults over 60 years of age.

243. The holder of the radiological authorization will promptly investigate the following:

1) any exposure for diagnostic purposes greater than or less than prescribed by the practitioner, leading to repeat the procedure, resulting in a significant increase in the established reference levels;

2) any malfunction or malfunction of equipment, accidents, errors that can lead to a patient's exposure significantly higher than prescribed.

244. In any of the above cases, the holder of the radiological authorization shall take measures to:

1) calculation and estimation of received doses by the patient;

2) indication and implementation of the necessary corrective measures, such as the reduction of the received doses by the patient, verification and adjustment of the used equipment, improvement of the applied procedures;

3) submitting a written report to the State Supervision Service of Public Health and the Agency on the causes of the incident within 24 hours from its production.

CHAPTER XXIII

MEDICAL RECORDS

245. The holder of the radiological authorization shall keep, for a period of 10 years, information about administered radiopharmaceuticals and activities to patients.

246. The clinical observation sheet for treated patients with radiopharmaceuticals is going to include the following mandatory data:

1) the administered radiopharmaceutical, route of administration, activity and date of administration;

2) the name of the person, who performed the control of the administered activity, the method of control and the data necessary for the identification of that radiopharmaceutical delivery;

3) the recording of made measurements at discharge.

CHAPTER XXIV

REQUIREMENTS FOR THE RECEPTION, STORAGE, STORAGE AND EVALUATION OF RADIOACTIVE WASTE

247. It are considered radioactive waste resulting from the practice of nuclear medicine:

1) closed radioactive sources, that were used for testing, quality control and equipment calibration;

2) consumed radionuclide generators;

3) prepared solutions, that are not used for diagnostic purposes;

4) very low radioactive liquid wastes from washing machines;

5) water immiscible liquids (scintillation fluids);

6) solid waste: pipettes, syringes used for radionuclides, filters;

7) very low radioactive solid waste: syringes, pipettes, glass, ampoules or other decontamination;

8) radioactive gases;

9) ashes from incineration;

10) waste arising from medical investigations.

248. The radioactive sources, including contaminated substances, materials and objects, that are resulting from nuclear medicine practice, may be released from regulatory requirements only if the concentration values ​​of specific activities / activity have reached values ​​less than or equal to the release levels .

249. The contaminated radioactive waste as well as sources of ionizing radiation, that can no longer be used in nuclear medicine practice, having a half-life of up to 100 days, may be collected and temporarily stored until they have reached lower values ​​or substantially equal to the release levels below the authorization requirements.

250. The solid waste is considered to be radioactive waste only if the activity of radionuclides emitting beta or gamma radiation exceeds 100kBq / kg (kiloBequerell / kg) and the activity of radionuclides emitting alpha radiation exceeds 10kBq / kg. The activity can be averaged for a quantity of waste not exceeding 10 kg. The solid wastes that are not considered as radioactive waste according to the mentioned criteria can be treated as usual laboratory waste.

251. Solutions containing scintillation fluids or other wastes containing organic solvents are considered to be radioactive waste only in cases where:

1) the specific concentration does not exceed 10 Bq / ml (Bequerel / milliliter) and there are no alpha emitting radionuclides;

2) the specific concentration does not exceed 100 Bq / ml and only 14C or 3H exists.

252. The limiting activity for the radioactive substances that can be discharged to the centralized sewerage system is set out in Annex no. 12 to these Regulations.

253. In the table in Annex no. 12 the ALImin values ​​for the most commonly used radionuclides are indicated. For the calculation of ALImin for other radionuclides will be used the dose limits and effective doses, employed on the incorporation unit according to the applicable radioprotection regulations.

254. The emission of gaseous radioactive substances into the environment must meet the relationship shown in Annex no. 13 to these Regulations.

255. In the table in annex no. 13 are exposed the CMA values ​​for the most commonly used radionuclides. For the calculation of air / CMA-derived concentration, for other radionuclides the dose limits and actual doses, employed on the inhalation incorporation unit will be used according to the applicable radiation protection regulations.

256. The closed radiation sources that cannot be handed over as regular waste shall be handed over to a authorized holder by the Agency for collection and disposal as radioactive waste or returned to the manufacturer.

257. The radiological authorization holder shall establish and maintain a primary and operative record of all types of radioactive waste - liquid, solid, gaseous. The record shall contain at least the following information:

1) the type of waste;

2) the radiopharmaceutical name / radionuclide designation;

3) the activity of waste at the date of storage;

4) the place of storage;

5) the activity of waste at the time of delivery / release to the environment;

6) the name of the authorized holder to whom the waste has been transferred, if applicable;

7) the name of the radiological security officer and the person, who handed over or removed the radioactive waste.

258. The receiving, storage and transfer of radioactive waste at storage shall be performed according to the provisions of the register in Annex no. 14 to these Regulations. For solid radioactive waste, columns 3 and 4 of the register are not filled in. In the case of the transmission of radioactive waste for centralized storage, column no. 11 of the register shall not be filled in. In the case of waste elimination after storage for disintegration as non-radioactive waste, columns 12 and 13 of the register are not filled in. In order to settle qualified waste as non-active after radioactive disintegration below the established limits, a report on the settlement of open radioactive sources shall be drawn up.

259. The holder of the radiological authorization shall establish and maintain a procedure for the collection, treatment and disposal of radioactive waste, arising from the practice of nuclear medicine, through which:

1)to ensure, that the activity and volume of any radioactive waste, arising from the radioactive sources for which it is responsible is kept as small as possible and that the waste is managed and managed;

2)to ensure, that the disposal of radioactive waste in public waste or sewerage systems is within the limits of the legislation in force;

3) to maintain their responsibilities over radioactive sources and waste until the forms for their transfer have been completed to a new authorized holder or have been handed over to a radioactive waste treatment, conditioning and storage facility.

260. The solid radioactive waste (cotton wool, gauze, used syringes, etc.) is collected in special equipped containers with replacement polyethylene bags, which at the end of the working day shall be handed over to the radioactive waste storage room. In case of simultaneous work with different radiopharmaceuticals, the waste will be collected in a separate container for each preparation. The radioactive waste with a mixed radionuclide content and with a half-life of less than 15 days is collected in a common container.

261. In the radioactive waste deposit, all solid wastes are sorted according to the radionuclide type and, after activity, placed in plastic bags and placed behind the radioprotection screen. Each bag shall be labeled with a badge containing the information on the radionuclide type, the date of placing on storage and the date of establishment as plain waste. Hereinafter the disintegrated waste is then managed along with regular waste and medical waste after mandatory dosimetric control and removal of the radioactive sign.

262. The used generators are stored in the radioactive waste repository and after disintegration they are transferred to the producing company or transferred as waste of used metallic stains.

263. The solid radioactive waste with half-life of radionuclides longer than 15 days will be transferred to specialized organizations for further processing and centralized storage.

CHAPTER XXV

RADIOLOGICAL EMERGENCIES, PLANNING AND PREPARING THE INTERVENTION. REPORTING ACCIDENTS

264. The radioprotection system for patients, personnel and persons in the population includes measures for the prevention and prevention of accidents and incidents with radiation, to ensure the appropriate actions of the personnel in case of accident / incident occurrence and the measures of exclusion and minimization the influence of radiation on the population during the liquidation of their consequences.

265. The radiological authorization holder is required to identify incidents / accidents that could affect occupational exposure and / or people in the population and to prepare an emergency procedure. The Emergency Procedure shall contain contingency plans for the radiological emergencies, that are identified and approved by the radiological authorization holder and coordinated with the State Health Service Agency and State Supervision Service.

266. The radiological authorization holder may contract the execution of the intervention in case of radiological emergency with an authorized intervention facility. The contract must clearly stipulate the responsibilities of both parties.

267. The events for which emergency plans are going to be drawn are  the fire, natural disasters (floods, earthquakes) and the following:

1) breaking the ampoule, containing the radioactive material or the 99Mo-99Tc generator;

2) loss of control over radioactive sources in the laboratory;

3) accidental contamination of areas;

4) breakage of radioactive or radioactive waste tanks;

5) administering to the patient an activity higher than that established by reference levels.

268. The Radiological Protection Intervention Plan for the Practice of Nuclear Medicine is going to be approved by the Administration of the Medical Institution and coordinated with the State Agency of Public Health Supervision and Service. The intervention plan must contain the following:

1) identifying potential accidents and other events that might occur during nuclear medicine practice and assessing the potential risks, associated with them;

2) identification of all authorized persons to participate in different phases of the intervention;

3) identification of all persons or organizations to be notified in the event of an incident, including the possibility of contacting them (telephone, fax, e-mail, addresses);

4) indications on how to recognize when a specific intervention plan is needed;

5) specific instructions for each identified emergency situation;

6) special instructions to be followed, when there is a risk of loss of life;

7) identification and acquisition of the means and equipment necessary for the intervention;

8) instructions for removing panic, forbidding to reach the radioactive sources by hand, remote manipulation, planning the next steps and their implementation, limiting to intervention measures that are consistent with the degree of experience and the competence available, the request for support or technical assistance.

269. The radioprotection and radiological safety officer is appointed by order of the radiological authorization holder and is responsible for the concrete implementation of the approved contingency plan.

270. The intervention exercises will be performed periodically for the implementation and application of the approved intervention plan. Records of exercises and earnings will be kept.

271. Prior to the intervention, the participants in the intervention are going to be warned about the health risks and to accept in writing the participation in the intervention.

272. The minimum required equipment for intervention in the case of radiological installations with radioactive sources is going to provide :

- the appropriate devices for handling the radiation sources;

2) the appropriate dosimetric control equipment;

3) the physical barriers and corresponding warning signs;

4) the lead screens for radioactive used sources;

5) the appropriate means for decontamination;    (6) suitable containers for storing radioactive sources or produced waste.

273. All designated persons to participate in the interventions are going to be nominated in writing, properly trained to understand the contents of the intervention plan and to use the means, equipment necessary for the intervention properly.

274. The knowledge of the assigned personnel to participate in the interventions is going to be periodically checked for a period of time set by the radiation protection rules.

275. The report on the occurrence of the radiological incident or incident shall be prepared by the radioprotection and radiological safety officer, with the participation of an accredited radiology expert and / or a physicist, as appropriate. The report should include:

1) the most detailed description of the incident or accident;

2) the identification of equipment and involved persons;

3) the causes of the accident, where they are known;

4) the taken measures to restore the situation;

5) the estimated or determined doses, received by the affected personnel;

6) the recommendations on the measures to be taken to prevent similar situations.

276. The holder of the radiological authorization has the obligation to announce within 10 days from the date of acquisition:

1) the Agency and the Ministry of Health about any incident or radiological accident through a report drawn up;

2) the Public Health Supervision Service and the Agency about any exposure of the patient at a higher dose than the prescribed one;

3) the local police body, the Ministry of Health, and the Agency about any loss or circumvention of radioactive sources through a report describing the circumstances in which they were lost or stolen, the quantities, extracted form and content and the risks that may arise.

Annex No.1

Annex No.1

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Composition and surface area

of in vivo nuclear medicine subdivisions

| No. n/o | Destination of the room | Surface, m2 (minimum) | Annotation |
| --- | --- | --- | --- |
| 1 | 2 | 3 | 4 |
| Assurance block with radiopharmaceuticals |  |  |  |
| 1. | Room for receiving and disposing of radiopharmaceuticals | 10 |  |
| 2. | Deposit of radiopharmaceutical preparation | 10 |  |
| 3. | Deposit of radioactive waste | 10 |  |
| 4. | Room of generators | 20 |  |
| 5. | Packaging room of radiopharmaceutical | 20 |  |
| 6. | Washing room | 10 | It can be combined with the packaging room or the procedure room |
| 7. | Main room of procedures | 20 |  |
| 8. | Secundary room of procedures | 10 | It can be used for the introduction of radiopharmaceuticals for ambulatory patients |
| 9. | WC for patients | 5 |  |
| 10. | Personal hygiene room | 5 |  |
| 11. | Storage room of auxiliary inventory | 3 | For cleaning only in the premises of the block |
| 12. | Room for storing contaminated clothing | 5 |  |
| in vivo radiopharmaceutical investigation block |  |  |  |
| 13 | Gamma-room cabinet | 20 |  |
| 14. | Control room of gamma-room | 10 |  |
| 15. | Cabinet with gamma-tomograph | 40 |  |
| 16. | Control room of gamma-tomography | 20 |  |
| 17. | in vivo radiometric room | 20 |  |
| 18. | Radiometric room of biomedies | 10 |  |
| 19. | Computer room | 10 |  |
| 20. | Waiting room | 10 |  |
| Block of general permises |  |  |  |
| 21. | Cabinet for Section Chief | 20 |  |
| 22. | Cabinet for doctors | 10 | 5,0 m2 for each |
| 23. | Cabinet for nurses | 10 | 3,0 m2 for each |
| 24. | Registry office | 10 |  |
| 25. | Section archive | 10 |  |
| 26. | Cabinet for medical physicist | 10 | 5,0 m2 for each |
| 27. | Cabinet for the programmer engineer with the workshop | 20 |  |
| 28. | Cabinet for superior medical care | 10 |  |
| 29. | Cabinet for sister-housewife with deposit | 20 |  |
| 30. | Storage room of auxiliary inventory | 3 |  |
| 31. | WC for personnel | 3 |  |

Annex No.2

Annex No.2

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Composition and surface of the center permises

of positron emission tomography (PET)

| No. n/o | Destination of the room | Surface, m2 (minimum) | Annotation |
| --- | --- | --- | --- |
| 1 | 2 | 3 | 4 |
| Assurance block with radiopharmaceuticals |  |  |  |
| 1. | The cyclotron bunker | 50 |  |
| 2. | Control Room | 20 |  |
| 3. | Technical room | 10 |  |
| 4. | Room of target preparation | 20 |  |
| 5. | Radiopharmaceutical synthesis room | 50 |  |
| 6. | Control Room of radiopharmaceutical quality | 20 |  |
| 7. | Gas balloon permises | 10 |  |
| 8. | Radiopharmaceutical deposit | 10 |  |
| 9. | Radioactive waste storage | 10 |  |
| 10. | Laundry | 10 | It can be merged with the procedure room |
| 11. | Room of procedures | 20 |  |
| 12. | Backup procedures room | 10 |  |
| 13. | Sanitary filterv | 10 |  |
| 14. | Storage room of the auxiliary inventory | 5 |  |
| 15. | Room for storing of contaminated clothing | 5 |  |
| Block of investigations with radiopharmaceuticals |  |  |  |
| 16. | Positron tomography cabinet | 50 |  |
| 17. | Control room of tomography | 20 |  |
| 18. | Technical room | 10 |  |
| 19. | Room for computerized processing of data | 20 |  |
| 20. | Room of dosimetry and radiometry | 20 |  |
| 21. | Chemical reagent storage, cyclotron, tomography and radiochemistry equipment | 20 |  |
| 22. | Waiting room | 10 |  |
| 23. | WC for patients | 5 |  |
| 24. | Storage room of auxiliary inventory | 5 | For cleaning only in the premises of the block |
| 25. | Workshop | 20 |  |
| Block of general rooms |  |  |  |
| 26. | Cabinet for chief | 20 |  |
| 27. | Cabinet for doctor | 10 | 5,0 m2 each for |
| 28. | Cabinet for assistants | 10 | 3,0 m2 each for |
| 29. | Registry office | 10 |  |
| 30. | Archive of the center | 10 |  |
| 31. | Cabinet for medical physicist | 10 | 5,0 m2 each for |
| 32. | Cabinet for radiochemist | 10 | 5,0 m2 each for |
| 33. | Cabinet for the radiophysicist engineer | 10 | 5,0 m2 each for |
| 34. | Cabinet for the electronics engineer | 10 | 5,0 m2 each for |
| 35. | Cabinet for sister-housewife with deposit | 20 |  |
| 36. | Storage room of auxiliary inventory | 5 |  |
| 37. | WC for personnel | 5 |  |
| 38. | Auxiliary room | 20 |  |

Annex No.3

Annex No.3

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Instruction record register for

radiation protection and safety techniques

| Date | Name and surname of the employee | Specialization and position (post) of the employee | Instruction: introductory, primary, repeated, unplanned | Name and surname of the instructor | Signature |  | Signature and post of the person, who allowed the activity |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | of employee | of instructor |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 |

Annex No.4

Annex No.4

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

List of minimum checks for

used equipment in nuclear medicine

| - Equipment | Acceptance / reference verification or periodic and / or routine verification (weekly, monthly, etc.) | Daily verification |
| --- | --- | --- |
| - Dose calibrator | Precision, accuracy, linearity, geometry, background | Reproductibilitate, fond |
| - Counting system | Scale / timer operation, energy calibration, energy resolution, sensitivity, counting accuracy, linearity to energy, background, activity linearity, geometry | Calibrare de energie, fond |
| - Gamma room (including SPECT) | PHA window setup, energy resolution, linearity in energy response, intrinsic uniformity, system uniformity, intrinsic linearity, intrinsic spatial resolution, intrinsic spatial resolution, counting performance, sensitivity, uniformity for the whole body, resolution for the whole body, pixel size, computer timing, rotation center, tomography uniformity, tomographic spatial resolution, total SPECT performance | Settings for PHA window, uniformity, sensitivity, background, rotation center, collimator |
| - Movie Processor | Wave level, contrast, resolution, speed, sensitivity | Wave level, contrast |
| - PET | Verification of calibration, uniformity, spatial resolution, scattering fraction, sensitivity, losses at counting speed, gliding in time coincidence measurement and energy threshold, mechanical motion of detector circles, verification of positioning seps, laser alignment, correction accuracy mitigation, accuracy to dead time corrections, accuracy to scaling corrections, accuracy to random coincidence corrections | Verification calibration |

Annex No.5

Annex No.5

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Record register of usage of received

radiopharmaceuticals in the nuclear medicine subdivision

for introduction into the patient's body

| No. n/o | Receipt |  |  |  |  | Use |  |  |  | Discount |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Date of entry | Number and date of delivery note | Number and date of passport release | Radiopharmaceutical container number | Passport activity, MBq | Whom is was delivered | date of release | release request number | Activity on release day, MBq | Settlement date | Settlement report number |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |

Annex No.6

Annex No.6

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Record register of usage

of short life generators

| No. n/o | Receipt |  |  |  |  | Use |  |  |  | Discount |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Date of entry | Number and date of delivery note | Number and date of passport release | Generator number | Passport activity, MBq | Whom is was delivered | Data eliberării | release request number | Activity on release day, MBq | Settlement date | Settlement report number |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |

Annex No.7

Anexa nr.7

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Registrul de pregătire a

soluţiilor de lucru  a radiofarmaceuticelor

| No. n/o | Initial solution (eluate) |  |  |  |  | Work solution (radiopharmaceuticals) |  |  |  | Signature of the medical assistant, who prepared the radiopharmaceutical solution |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Date of preparation of the eluate | Generator number | Term of validity (according to the passport) | Volume of the eluate, ml | Activity of the eluate, МBq | Name of radiopharmaceuticals | Volume of the solution, ml | Activity of the solution, MBq | Data întroducerii pacientului conform registrului |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 |

Annex No.8

Annex No.8

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Administration registry

of radiopharmaceuticals for patients

| No. n/o | Date of administration | Name, surname of the patient | Medical card number or ambulatory card numberde ambulator | The section, that directed the patient for investigation | Other institutions, that directed the patient for investigation | Introduced radiopharmaceutical preparation |  |  | Signatures |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | Name | Volume, ml | Activity of the solution, МBq | Radiologist, who sent the patient to the investigation | The medical assistant, who introduced the preparation |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 |

Annex No.9

Anexa nr.9

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Gamma equivalent dose rates for classification of areas

| Dose rate | Value for the controlled area | Value for the supervised area | Value for the unsupervised area |
| --- | --- | --- | --- |
| 1. Instantly, median one minute (IDR) | >2000μSv/h | >7,5 μSv/h | <7,5 μSv/h |
| 2. Mediated for 8 hours, taking into account the use and charging, for the worst case (TDR) | >7,5 μSv/h | >0,5 μSv/h | <0,5 μSv/h |
| 3. Mediated for 2000 hours, taking into account the use and charging, for the worst case (TADR) | >3 μSv/h | >0,15 μSv/h | <0,15 μSv/h |

Annex No.10

Annex No.10

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Nivelurile de referinţă a activităţii radiofarmaceuticilor în timpul unei

investigaţii de medicină nucleară pentru un pacient adult-standard

| Investigation | Radionuclide |  | Chemical form | Maximum activity for diagnosis (МBq) |
| --- | --- | --- | --- | --- |
| 1. Bones |  |  |  |  |
| 1.1. Radioscintigraphy of bones | Tc–99m |  | Phosphonates and phosphate compounds | 600 |
| 1.2. Bone Radioscintigraphy SPECT (Single Photon Emission Computerized Tomography) | Tc–99m |  | Phosphonates and phosphate compounds | 800 |
| 1.3. Scintigraphy of the bone marrow | Tc–99m |  | Marked colloidal solution | 400 |
| 2. Brain |  |  |  |  |
| 2.1. Scintigraphy of the brain (in static) | Tc–99m |  | TcO4- , DTPA, gluconate and glucoheptonate | 500 |
| 2.2. Scintigraphy of the brain (SPECT) | Tc–99m |  | TcO-4, pertechnetate DTPA, gluconate and glucoheptonate Ekzametazin (Ceretec) | 800 800 500 |
| 2.3. Investigation of cerebral blood flow | Tc–99m |  | Hexamethyl propylene aminooksim (HM-PAO), Ekzametazin | 500 |
| 2.4. Cisternography | In–111 |  | DTPA | 40 |
| 3. Lacrimal gland |  |  |  |  |
| 3.1. Drainage | Xe–133 Tc–99m |  | TcO41, pertechnetate labeled Colloidal solution, pertechnetate | 4 |
| 4. Thyroid gland |  |  |  |  |
| 4.1. Scintigrafia glandei tiroide | Tc–99m I–123 |  | TcO-4, pertechnetate I- | 100 20 |
| 4.2. Diagnosing of the metastases (after removal) | I–131 |  | I- | 400 |
| 4.3. Scintigraphy of the parathyroid gland | Тl–201 Tc–99m |  | Tl+ chloride Tetrofosmin (Myoview) | 80 400 |
| 5. Lungs |  |  |  |  |
| 5.1. Scintigraphy of pulmonary ventilation | Kr–81m Tc–99m |  | Gaze DTPAaerosol (Diethylenetriaminepenta-acetic acid) | 600 800 |
| 5.2. Scintigraphy of the pulmonary perfusion | Tc–99m |  | Human albumin (in macro aggregates (MAA) or in microspheres) | 100 |
| 5.3. Scintigraphy of the pulmonary perfusion (planar) | Хе–133 |  | Isotonic solution | 200 |
| 5.4. Scintigraphy of the pulmonary perfusion (SPECT) | Tc–99m |  | МАА(macroagregaded albumin) | 200 |
| 6. Liver and spleen |  |  |  |  |
| 6.1. Scintigraphy of the liver and spleen | Tc–99m |  | Marked colloidal solution | 80 |
| 6.2. Scintigraphy of the biliary system function | Tc–99m |  | Iminodiacetate and their equivalents | 150 |
| 6.3. Scintigraphy of the spleen | Tc–99m |  | Denatured erythrocytes | 100 |
| 6.4. Scintigraphy of the liver (SPECT) | Tc–99m |  | Marked colloidal solution | 200 |
| 7. Circulatory system |  |  |  |  |
| 7.1. Myocardial scintigraphy |  | Tc–99m | Phosphonates and phosphate compounds | 600 |
| 7.2. Myocardial diagnosis (SPECT) |  | Tc–99m Tl–201 Tc–99m Tc–99m | Isonitrile Tl+ chloride Phosphonates and phosphate compounds Tetrofosmin (Myoview) | 300 100 800 800 |
| 8. Stomach, digestive tract |  |  |  |  |
| 8.1. Stomach scintigraphy (salivary gland) |  | Tc–99m | TcO-4 | 40 |
| 8.2. Meckel diverticulus scintigraphy |  | Tc–99m | TcO-4 | 400 |
| 8.3.Scintigrafia hemoragiei în tractul gastrointestinal |  | Tc–99m | Marked colloidal solution Marked normal erythrocytes | 400 |
| 8.4. Scintigraphy of permeability and gastroesophageal reflex |  | Tc–99m | Marked colloidal solution | 400 |
| 8.5. Scintigraphy of gastric emptying |  | Tc–99m In–111 In–113m | Unabsorbed compounds | 12 |
| 9. Kidney, urinary system and adrenal glands |  |  |  |  |
| 9.1. Scintigraphy of the kidneys (in static) |  | Tc–99m | DTPA, gluconate and glucoheptonate | 350 |
| 9.2. Scintigraphy of the kidneys (in dynamics) |  | Tc–99m I–123 | Macroaggregated globulin О- iodohipurat | 100 20 |
| 9.3. Scintigraphy of the adrenals |  | Se–75 | Selenocolesterin | 8 |
| 10. Other organs |  |  |  |  |
| 9.4. Scintigraphy of the neuroectodermal tumor |  | I–123 I–131 | Meta-iodo-benzyl-guanidine | 400 20 |
| 9.5. Scintigraphy of the lymph node |  | Tc–99m | Marked colloidal solution Nanokol | 80 140 |
| 9.6. Scintigraphy of the abscess |  | Tc–99m In–111 | Labeled white blood cells, treated with eksametazin Marked leukocytes | 400 20 |
| 9.7. Scintigraphy of the thrombus |  | In–In | Platelets marked | 20 |

Annex No.11

Anexa nr.11

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Recommendations on interruption periods of breastfeeding

to ensure the effective dose of the baby is less than 1 mSv

| No. n/o | Radiopharmaceutical | Activity | Instructions |
| --- | --- | --- | --- |
| 1 | 2 | 3 | 4 |
| 1. | 51Cr-EDTA | * | Interruption of breastfeeding is not usually necessary |
| 2. | 99mTc DISIDA | * | Interruption of breastfeeding is not usually necessary |
| 3. | 99mTc DMSA | * | Interruption of breastfeeding is not usually necessary |
| 4. | 99mTc DTPA | * | Interruption of breastfeeding is not usually necessary |
| 5. | 99mTc diphosphonate | * | Interruption of breastfeeding is not usually necessary |
| 6. | 99mTc glucoheptonate | * | Interruption of breastfeeding is not usually necessary |
| 7. | 99mTc gluconate | * | Interruption of breastfeeding is not usually necessary |
| 8. | 99mTc HMPAO | * | Interruption of breastfeeding is not usually necessary |
| 9. | 99mTc MIBI | * | Interruption of breastfeeding is not usually necessary |
| 10. | 99mTc colloidal sulfide | * | Interruption of breastfeeding is not usually necessary |
| 11. | 111In leukocytes | 20 MBq | Interruption of breastfeeding is required if the activity is greater than that in column 3 |
| 12. | 201Tl clorid | 80 MBq | Interruption of breastfeeding is required if the activity is greater than that in column 3 |
| 13. | 99mTc MAG3 | 100 MBq | Interruption of breastfeeding is required if the activity is greater than that in column 3 |
| 14. | 99mTc MAA | 100 MBq | Interruption of breastfeeding is required for 13 hours |
| 15. | 99mTc pertechnetate | 800 MBq | Interruption of breastfeeding is required for 48 hours |
| 16. | 99mTc pertechnetate | 80 MBq | Interruption of breastfeeding is required for 24 hours |
| 17. | 99mTc erythrocytes |  | Interruption of breastfeeding in function of activity concentration in milk |
| 18. | 99mTc aerosol |  | Interruption of breastfeeding in function of activity concentration in milk |
| 19. | 99mTc MAG3 | >100MBq | Interruption of breastfeeding in function of activity concentration in milk |
| 20. | 99mTc microspheres |  | Interruption of breastfeeding in function of activity concentration in milk |
| 21. | 99mTc pyrophosphate |  | Interruption of breastfeeding in function of activity concentration in milk |
| 22. | 123I iodide |  | Interruption of breastfeeding in function of activity concentration in milk |
| 23. | 123I MIBG |  | Interruption of breastfeeding in function of activity concentration in milk |
| 24. | 123I hipuran |  | Interruption of breastfeeding in function of activity concentration in milk |
| 25. | 32P sodium phosphonate |  | Interruption of breastfeeding in function of activity concentration in milk |
| 26. | 67Ga citrate |  | Interruption of breastfeeding in function of activity concentration in milk |
| 27. | 123I HSA |  | Interruption of breastfeeding in function of activity concentration in milk |
| 28. | 131I iodide |  | Interruption of breastfeeding in function of activity concentration in milk |

Annex No.12

Annex No.12

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Limit activity for radioactive substances,

that can be discharged into the centralized sewerage system

1. The maximum total activity, that can be released into centralized sewerage through a single spill fulfills the condition:

2. The maximum total activity that can be released into centralized sewerage which in one month fulfills the condition:

where  − represents the activity of the k  radionuclide ;

− is the smallest calculated value for the annual incorporation limit for the k radionuclide.

Table

Values  for the most often used

radionuclides in nuclear medicine

| Radionuclide | (Bq) |
| --- | --- |
| 3H | 1.109 |
| 14C | 3.107 |
| 18F | 2.108 |
| 32P | 6.106 |
| 51Cr | 5.108 |
| 57Co | 2.107 |
| 60Co | 7.105 |
| 67Ga | 7.107 |
| 68Ga | 2.108 |
| 89Sr | 3.106 |
| 90Y | 7.106 |
| 99Mo | 2.107 |
| 99mTc | 7.108 |
| 111In | 7.107 |
| 123I | 1.108 |
| 124I | 2.106 |
| 125I | 1.106 |
| 131I | 9.105 |
| 137Cs | 2.106 |
| 153Sm | 3.107 |
| 186Re | 1.107 |
| 201Tl | 2.108 |

Annex No.13

Annex No.13

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Emission of gaseous radioactive substances into the environment

Emission of gaseous radioactive substances into the environment must meet the relationship shown:

where:

− is the concentration of k radionuclide activity;

− is the maximum permitted concentration for k radionuclide and is equal to 1% of the derived concentration k radionuclide in air.

Table

CMA values for the most used radionuclides

in nuclear medicine

| radionuclide | CMA (Bq/m3) |
| --- | --- |
| 3H | 3.103 |
| 14C | 1.102 |
| 18F | 5.102 |
| 32P | 3.101 |
| 51Cr | 2.103 |
| 57Co | 9.101 |
| 60Co | 3.100 |
| 67Ga | 3.102 |
| 68Ga | 1.103 |
| 89Sr | 1.101 |
| 90Y | 5.101 |
| 99Mo | 8.101 |
| 99mTc | 3.103 |
| 111In | 3.102 |
| 123I | 4.102 |
| 124I | 7.100 |
| 125I | 6.100 |
| 131I | 4.100 |
| 137Cs | 1.101 |
| 153Sm | 1.102 |
| 186Re | 7.101 |
| 201Tl | 1.103 |

Annex No.14

Annex No.14

to the Sanitary Regulations on Radiorotection

and Radiological Safety

in Nuclear Medicine Practices

Radioactive waste records register

| No. n/o | Date of entry | Liquid waste |  | Composition of radionuclides | Summary activity, MBq | From whom it was received |  | discount |  |  | Transmission to storage |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | pH solution | Volume, ml |  |  | Name, surname | Signature, who surrendered waste | Settlement date | Settlement release Number | Storage date after storage for disintegration | Date of transmission | Passport number for handing over the radioactive waste |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 |